## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E1413-19
ANSI/ASTM E1413-19
Standard Practice for Separation of Ignitable Liquid Residues from Fire Debris Samples by Dynamic Headspace Concentration onto an Adsorbent Tube
Discipline: | Fire Debris & Explosives
Checklist version: | V1.1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Apparatus | 6 | Section Title
Positive Applied Pressure-Based Sampling System | 6.1 | Informational Text | A system capable of increasing the pressure inside the sample container by applying dry inert gas at the sample container inlet, thereby pushing the headspace vapors toward and through the outlet.
Positive Applied Pressure-Based Sampling System | 6.1.1 | Requirement | Sample Pressurization Device A system capable of delivering pressurized dry inert gas at up to approximately 276 kPa (40 psi) to the inlet.
Positive Applied Pressure-Based Sampling System | 6.1.1.1 | Requirement | A needle valve capable of fine control of the flow rate of the dry inert gas, at up to approximately 400 mL/min.
Positive Applied Pressure-Based Sampling System | 6.1.2 | Requirement | A filtering tube, consisting of an adsorbent tube as described in 7.5, attached to the sample container inlet, for the purpose of filtering incoming dry inert gas.
Positive Applied Pressure-Based Sampling System | 6.1.3 | Requirement | A flow meter capable of measuring the flow of the dry inert gas through the end of the adsorption tube at the rate of approximately 40 to 400 mL/min.
Positive Applied Pressure-Based Sampling System | 6.1.4 | Requirement | Design and use the sampling system in a manner that prevents the loss of volatile compounds from the container and contamination of the sample from the apparatus itself or the laboratory environment.
Negative Applied Pressure-Based Sampling System | 6.2 | Informational Text | A system capable of creating partial vacuum inside the sample container by connecting a vacuum system to the outlet, thereby pulling the headspace vapors toward and through the outlet.
Negative Applied Pressure-Based Sampling System | 6.2.1 | Requirement | A vacuum system capable of pulling between approximately 40 and 400 mL/min at the outlet.
Negative Applied Pressure-Based Sampling System | 6.2.2 | Requirement | A filtering tube, consisting of an adsorbent tube as described in 7.5, attached to the sample container inlet, for the purpose of filtering incoming air.
Negative Applied Pressure-Based Sampling System | 6.2.3 | Requirement | A flow meter capable of measuring the flow of filtered air and headspace vapors through the outlet at the rate of approximately 40 to 400 mL/min.
Negative Applied Pressure-Based Sampling System | 6.2.4 | Requirement | Design and use the sampling system in a manner that prevents the loss of volatile compounds from the container and contamination of the sample from the apparatus itself or the laboratory environment.
Apparatus | 6.3 | Requirement | Heating System An oven large enough to accommodate the sample container, or a heating mantel designed to fit the sample container, capable of maintaining the required temperature range uniformly throughout.
Apparatus | 6.4 | Requirement | Temperature Measuring Device A thermometer or thermocouple capable of measuring the temperature of the heating system in the required range of operation, to within approximately 5°C.
Apparatus | 6.5 | Requirement | Thermal Desorption System A system capable of desorbing trapped volatile compounds from an adsorbent tube by means of elevated temperature, refocusing them on a cold-trap, and subsequently introducing them to a capillary GC column by flash heating. The desorption apparatus is directly coupled to a GC-MS.
Apparatus | 6.6 | Requirement | Puncturing Device A device, such as a nail, that is capable of puncturing small holes in the lid of the sample container. The size of the holes is such that the adsorbent tubes can be introduced.
Reagents And Materials | 7 | Section Title
Reagents And Materials | 7.1 | Requirement | Purity of Reagents Use reagent grade or better chemicals in all tests. Unless otherwise indicated, it is intended that all reagents conform to the specifications of the Committee on Analytical Reagents of the American Chemical Society where such specifications are available. Other grades may be used, provided it is first ascertained that the reagent is of sufficiently high purity through evaluation of appropriate blank(s) to permit its use without lessening the accuracy of the determination.
Reagents And Materials | 7.2 | Requirement | Filter Permeable material such as silanized glass wool, wire mesh or cotton held in place with a frit.
Reagents And Materials | 7.3 | Requirement | Solid Adsorbent Activated carbon, Tenax TA or equivalent.
Reagents And Materials | 7.4 | Requirement | Sampling Tubes Glass Pasteur pipettes or equivalent glass tubes for solvent elution, and stainless steel or glass tubes for thermal desorption.
Reagents And Materials | 7.5 | Requirement | Adsorbent Tubes Sampling tubes packed with a solid adsorbent.
Reagents And Materials | 7.5.1 | Informational Text | Pre-packed adsorbent tubes are commercially available. Empty sampling tubes that can be packed by the purchaser are also available.
Reagents And Materials | 7.5.1.1 | Informational Text | Activated carbon tubes and equivalent for solvent elution can be made by inserting a filter into the bottom of a sampling tube (approximately 5-mm diameter), then adding 2.5 cm to 5 cm of activated carbon or equivalent, which is held in place with a second filter.
Reagents And Materials | 7.5.1.2 | Informational Text | Tenax TA tubes and equivalent for thermal desorption can be made by inserting a filter into one end of a sampling tube, adding approximately 80 mg of pre-conditioned Tenax TA or equivalent, and then packing tightly with a second filter. NOTE 1: Tenax TA or equivalent is conditioned by heating, in accordance with instrument manufacturer or supplier instruction.
Reagents And Materials | 7.6 | Requirement | Screw cap or crimp-top glass vials with polytetrafluoroethylene (PTFE) lined seals.
Reagents And Materials | 7.7 | Requirement | Tape, rubber sleeve stopper, or equivalent.
Reagents And Materials | 7.8 | Requirement | Elution Solvent Suitable elution solvents include carbon disulfide, n-pentane, dichloromethane and diethyl ether.
Sample Preparation | 8 | Section Title
Sample Preparation | 8.1 | Requirement | Observe the appropriate procedures for handling and documentation of all submitted samples (see Guide E1459 and Practice E1492).
Sample Preparation | 8.2 | Requirement | Prepare the fire debris sample container for sampling by dynamic headspace concentration.
Sample Preparation | 8.2.1 | Informational Text | The sampling system is designed so that the headspace sample can be extracted from the container in which it was received at the laboratory.
Sample Preparation | 8.2.1.1 | Informational Text | Alternatively, the fire debris sample, or a portion of it, can be transferred to a clean, rigid sample container suitable for dynamic headspace sampling, such as a metal can with a friction-seal lid.
Sample Preparation | 8.2.1.2 | Requirement | Verify the cleanliness of the transfer container prior to sample transfer. Cleanliness is determined by means of analysis of a dynamic headspace concentration sample obtained from the empty transfer container using the same conditions as will be used for the questioned sample.
Sample Preparation | 8.2.1.3 | Requirement | Allow the transfer container with sample to equilibrate for at least one hour before sampling.
Sample Preparation | 8.2.2 | Requirement | Create two holes in the lid of the rigid sample container using a puncturing device, such that the adsorbent tubes can be introduced.
Sample Preparation | 8.2.2.1 | Requirement | Seal the holes with tape, rubber sleeve stopper, or equivalent. NOTE 2: Cans that are designed for fire debris samples, with a hole pre-fitted with a rubber sleeve stopper, are commercially available.
Adsorption Procedure | 9 | Section Title
Adsorption Procedure | 9.1 | Informational Text | A laboratory typically uses either a positive applied pressure-based system or a negative applied pressure-based system. Either type of system can be used with the adsorbent tubes described in 7.5.
Positive Applied Pressure-Based System | 9.2 | Section Title
Positive Applied Pressure-Based System | 9.2.1 | Requirement | Place the inlet inert gas filtering adsorption tube and the outlet sample adsorption tube through the holes in the lid of the sample container by penetrating the tape, rubber sleeve stoppers, or equivalent covering the holes.
Positive Applied Pressure-Based System | 9.2.1.1 | Requirement | Use a new inlet inert gas filtering adsorption tube for each sample container.
Positive Applied Pressure-Based System | 9.2.2 | Requirement | Connect the sample pressurization device to the inlet inert gas filtering adsorption tube. Then connect the exhaust tube to the outlet sample adsorption tube. Direct the system exhaust to a lab fume hood or equivalent.
Positive Applied Pressure-Based System | 9.2.3 | Requirement | Deliver the dry inert gas to the sample container using the sample pressurization device and adjust the flow rate.
Positive Applied Pressure-Based System | 9.2.3.1 | Informational Text | A suitable flow rate is determined using the method verification as outlined in Section 11. Typical flow rates are within the range of approximately 40 and 400 mL/min.
Positive Applied Pressure-Based System | 9.2.4 | Requirement | Heat the sample container, when necessary, during the adsorption procedure.
Positive Applied Pressure-Based System | 9.2.4.1 | Recommendation | Heating is not recommended if there are other evidentiary considerations such as preservation of DNA or latent fingerprints.
Positive Applied Pressure-Based System | 9.2.4.2 | Requirement | Place the sample container into the heating system at a suitable heating temperature, based upon parameters determined during the method verification as outlined in Section 11. NOTE 3: Sampling at room temperature is suitable for the extraction of the more volatile compounds, while heating increases the concentration of the less volatile compounds in the headspace.
Positive Applied Pressure-Based System | 9.2.4.3 | Requirement | Heat the container of fire debris at the selected temperature in order to release any ignitable liquid residues as a vapor into the headspace.
Positive Applied Pressure-Based System | 9.2.4.4 | Informational Text | Heating the adsorbent tube while extracting wet samples can prevent condensation from interfering with the adsorption process.
Positive Applied Pressure-Based System | 9.2.5 | Requirement | Continue the adsorption procedure for the period of time determined during the method verification as outlined in Section 11.
Positive Applied Pressure-Based System | 9.2.5.1 | Informational Text | Adsorption times can vary from minutes to hours, depending on the selected heating temperature, the size of the container, the concentration of ignitable liquid residues present in the sample, the type and amount of sample material in the container, and the selected adsorbent type and desorption procedure (solvent elution or thermal desorption).
Positive Applied Pressure-Based System | 9.2.5.2 | Informational Text | An indication of the ignitable liquid residue concentration in a questioned sample can be obtained from a preanalysis using an alternative method such as static headspace sampling (Practice E1388) at room temperature, or elevated temperature, and subsequent analysis with gas chromatography-flame ionization detection (GC-FID) or GC-MS.
Positive Applied Pressure-Based System | 9.2.6 | Requirement | Upon completion of the adsorption procedure, turn off the flow of dry inert gas, remove the inlet inert gas filtering adsorption tube and the outlet sample adsorption tube, and then seal the openings in the container lid with tape, rubber sleeve stoppers, or equivalent.
Positive Applied Pressure-Based System | 9.2.6.1 | Requirement | If an adsorption tube is to be thermally desorbed, recap it immediately.
Positive Applied Pressure-Based System | 9.2.6.2 | Requirement | If an adsorption tube is to be desorbed by solvent elution and it was heated, allow the adsorption tube to cool to room temperature prior to performing the desorption procedure.
Negative Applied Pressure-Based System | 9.3 | Section Title
Negative Applied Pressure-Based System | 9.3.1 | Requirement | Place the inlet air filtering adsorption tube and the outlet sample adsorption tube through the holes in the lid of the sample container by penetrating the tape, rubber sleeve stoppers, or equivalent covering the holes.
Negative Applied Pressure-Based System | 9.3.1.1 | Requirement | Use a new inlet air filtering adsorption tube for each sample container.
Negative Applied Pressure-Based System | 9.3.2 | Requirement | Apply a vacuum to the outlet sample adsorption tube using the vacuum system, and direct the system exhaust to a lab fume hood or equivalent. Adjust the applied vacuum to achieve the selected flow rate.
Negative Applied Pressure-Based System | 9.3.2.1 | Informational Text | A suitable flow rate is determined during the method verification as outlined in Section 11. Typical flow rates are within the range of approximately 40 and 400 mL/min.
Negative Applied Pressure-Based System | 9.3.3 | Requirement | Heat the sample container, when necessary, during the adsorption procedure.
Negative Applied Pressure-Based System | 9.3.3.1 | Recommendation | Heating is not recommended if there are other evidentiary considerations such as preservation of DNA or latent fingerprints.
Negative Applied Pressure-Based System | 9.3.3.2 | Requirement | Place the sample container into the heating system at a suitable heating temperature, based upon parameters determined during the method verification as outlined in Section 11. NOTE 4: Sampling at room temperature is suitable for the extraction of the more volatile compounds, while heating increases the concentration of the less volatile compounds in the headspace.
Negative Applied Pressure-Based System | 9.3.3.3 | Requirement | Heat the container of fire debris at the selected temperature in order to release any ignitable liquid residues as a vapor into the headspace.
Negative Applied Pressure-Based System | 9.3.3.4 | Informational Text | Heating the adsorbent tube while extracting wet samples can prevent condensation from interfering with the adsorption process.
Negative Applied Pressure-Based System | 9.3.4 | Requirement | Continue the adsorption procedure for the period of time determined during the method verification as outlined in Section 11.
Negative Applied Pressure-Based System | 9.3.4.1 | Informational Text | Adsorption times can vary from minutes to hours, depending on the selected heating temperature, the size of the container, the concentration of ignitable liquid residues present in the sample, the type and amount of sample material in the container, and the selected adsorbent type and desorption procedure (solvent elution or thermal desorption).
Negative Applied Pressure-Based System | 9.3.4.2 | Informational Text | An indication of the ignitable liquid residue concentration in a questioned sample can be obtained from a preanalysis using an alternative method such as static headspace sampling (Practice E1388) at room temperature, or elevated temperature, and subsequent analysis with gas chromatography-flame ionization detection (GC-FID) or GC-MS.
Negative Applied Pressure-Based System | 9.3.5 | Requirement | Upon completion of the adsorption procedure, turn off the vacuum system, remove the inlet air filtering adsorption tube and the outlet sample adsorption tube, and then seal the openings in the container lid with tape, rubber sleeve stoppers, or equivalent.
Negative Applied Pressure-Based System | 9.3.5.1 | Requirement | If an adsorption tube is to be thermally desorbed, recap it immediately.
Negative Applied Pressure-Based System | 9.3.5.2 | Requirement | If an adsorption tube is to be desorbed by solvent elution and it was heated, allow the adsorption tube to cool to room temperature prior to performing the desorption procedure.
Desorption Procedure | 10 | Section Title
Desorption Procedure | 10.1 | Informational Text | Solvent elution is typically used with activated carbon or equivalent adsorbents:
Desorption Procedure | 10.1.1 | Requirement | Follow the safety precautions described in the Safety Data Sheet (SDS) of the elution solvent that is used.
Desorption Procedure | 10.1.2 | Requirement | Pass approximately 1 mL of elution solvent through the adsorbent tube (containing activated carbon or the equivalent) into a vial. If possible, allow the solvent to remain in contact with the adsorbent material for approximately 5 minutes before expelling. An aspirator bulb can be used if necessary to push the solvent through the tube.
Desorption Procedure | 10.1.2.1 | Requirement | Alternatively, completely remove the adsorbent material from the tube and agitate up to approximately 5 minutes in a container with elution solvent. Transfer the solvent extract into a vial and cap the vial in preparation for analysis.
Desorption Procedure | 10.1.3 | Informational Text | If necessary, the solvent extract can be concentrated by evaporation under a stream of dry inert gas, dry nitrogen, or dry air that is filtered to remove hydrocarbons. (Warning—Evaporation can result in loss of some volatile compounds.)
Desorption Procedure | 10.1.4 | Informational Text | Refer to Practice E2451 for short-term and long-term storage of ignitable liquid extracts.
Desorption Procedure | 10.2 | Informational Text | Thermal desorption is typically used with Tenax TA or equivalent adsorbents:
Desorption Procedure | 10.2.1 | Requirement | Transfer the adsorbent tube to the thermal desorption system for desorption and subsequent GC-MS analysis.
Desorption Procedure | 10.2.1.1 | Informational Text | In the thermal desorption system, the volatile compounds trapped on the adsorbent tube are desorbed and refocused onto a cold trap prior to introduction into the GC-MS.
Quality Assurance | 11 | Section Title
Quality Assurance | 11.1 | Requirement | Verify the practice of dynamic headspace concentration onto an adsorbent tube prior to use on questioned samples utilizing Test Method E1618.
Quality Assurance | 11.1.1 | Requirement | Use a selection of ignitable liquids that together cover the entire range of compounds to be identified with this practice, a range of concentrations of those ignitable liquids, and ignitable liquids on matrices such as wood and carpet created to simulate common questioned samples.
Quality Assurance | 11.1.2 | Requirement | Determine the maximum flow rate of dry inert gas or air that can be swept through the outlet sample adsorption tube without analyte breakthrough, for common ignitable liquid classes described in Test Method E1618. Do not exceed this volume in practice.
Quality Assurance | 11.1.2.1 | Requirement | Prepare sample containers with a quantity of analyte mixture or ignitable liquid that is approximately ten times the upper limit of detection. Perform the adsorption procedure from each container in the same manner as questioned samples (as described in Sections 8 and 9), but place two adsorbent tubes in series, and use a different flow rate for each container.
Quality Assurance | 11.1.2.2 | Requirement | Analyze the secondary adsorbent tubes for the presence of ignitable liquid compounds, using the same desorption parameters as will be used for questioned samples (as described in Section 10), and using the same analysis parameters used for questioned samples in accordance with Test Method E1618.
Quality Assurance | 11.1.2.3 | Informational Text | Analyte breakthrough during adsorption is possible when using dynamic headspace concentration. Loss of some analytes can occur if the flow rate is too high.
Quality Assurance | 11.1.3 | Requirement | Determine the heating system temperature and the time period for heating of the sample container.
Quality Assurance | 11.1.3.1 | Informational Text | Heating is not required to recover lower boiling compounds up to C10, but can improve recovery.
Quality Assurance | 11.1.3.2 | Informational Text | Heating is required to recover higher boiling compounds, at temperatures typically between 60°C and 90°C for compounds up to C16, and as high as 130°C for compounds above C20.
Quality Assurance | 11.1.3.3 | Informational Text | A temperature below 90°C is most appropriate for samples that contain moisture or that are wet.
Quality Assurance | 11.1.3.4 | Informational Text | The heating time can vary from minutes up to several hours, depending on the selected temperature, the size of the container, and the type and amount of sample matrix in the container. (1) The required heating time generally increases with increased container volume or increased amount of sample matrix occupying that volume, or both.
Quality Assurance | 11.2 | Requirement | Store adsorbent material and capped adsorption tubes in a clean environment at all times to avoid adsorption of interference compounds.
Quality Assurance | 11.3 | Recommendation | Sample containers subjected to dynamic headspace concentration could be unsuitable for re-sampling. A portion of the concentrated headspace sample should be preserved for potential future analysis, if possible and if required, in accordance with Practice E2451.
Quality Control | 12 | Section Title
Quality Control | 12.1 | Requirement | Check the purity of each new bottle of the elution solvent before use, by evaporating a portion of the solvent to at least 50 % its original volume, or, if concentration of a solvent extract is performed, by evaporating at least twice the volume used in the analysis. Evaporation is performed as described in 10.1.3. Analyze the evaporated solvent using the same conditions as used for questioned samples and in accordance with Test Method E1618.
Quality Control | 12.2 | Requirement | Check each new, cleaned, or reconditioned batch of adsorbent tubes for contaminants before use by desorbing, at a minimum, at least one of the tubes using the same desorption procedure as used for questioned samples (as described in Section 10). Analyze these adsorbent tubes, or the samples prepared from these adsorbent tubes, using the same conditions as used for questioned samples and in accordance with Test Method E1618.
Quality Control | 12.2.1 | Requirement | Activated carbon or equivalent is considered to be disposable, whereas Tenax TA or equivalent is reusable. Check reusable adsorbents batch-wise, at a minimum, after being cleaned or reconditioned, prior to being re-used.
Quality Control | 12.2.1.1 | Informational Text | The lifetime of reusable adsorbents is usually specified by the suppliers and thermal desorption instrument vendors, but can be exceeded in practice, if their adsorption and desorption efficiencies are regularly checked as described in 12.4.
Quality Control | 12.3 | Requirement | Check the cleanliness of the sampling system and desorption procedure for contaminants prior to processing a batch of questioned samples, which can include preparing an empty container as a negative control. Prepare and sample the empty container in the same manner as questioned samples (as described in Sections 8 and 9). Desorb the adsorbent tube using the same desorption process and conditions as will be used for questioned samples (as described in Section 10), and analyze using the same conditions used for questioned samples and in accordance with Test Method E1618.
Quality Control | 12.3.1 | Informational Text | A negative control is a clean, empty sample container similar to that used for questioned samples.
Quality Control | 12.4 | Requirement | Check the adsorption and desorption efficiency of the system near the beginning and end of every sequence of questioned samples, or at a frequency established during method verification, by preparing a positive control containing a known volume of ignitable liquid, or of a standard test mixture in accordance with Test Method E1618 that spans a range of boiling points. Prepare and sample the container in the same manner as questioned samples (as described in Sections 8 and 9). Desorb the adsorbent tube using the same desorption process and conditions as will be used for the questioned samples (as described in Section 10), and analyze using the same conditions used for questioned samples in accordance with Test Method E1618.
Quality Control | 12.4.1 | Informational Text | The positive control is used to verify that the adsorption and desorption efficiency of the system is within the required specifications from the verification studies of the practice. The adsorption and desorption efficiency is determined by evaluating the recoveries of selected compounds of interest using peak area or peak height.
Quality Control | 12.4.2 | Informational Text | The positive control includes a periodic check of the adsorption and desorption capabilities of the individual or batch of adsorbent tubes in use.
Quality Control | 12.5 | Requirement | Check empty sampling tubes as in accordance with laboratory policy to demonstrate the cleanliness of the desorption process and the GC-MS. Desorb the empty sampling tubes using the same desorption process and conditions as used for questioned samples (as described in Section 10), and analyze using the same conditions used for questioned samples in accordance with Test Method E1618.
Quality Control | 12.6 | Requirement | If an internal standard is routinely added to questioned samples and reference samples, then also include it in the prepared blank.
Conditioning and Regeneration: | 12.7 | Section Title
Conditioning and Regeneration: | 12.7.1 | Requirement | Condition Tenax TA or equivalent adsorbents at elevated temperatures before initial use according to the laboratory or manufacturer’s instructions.
Conditioning and Regeneration: | 12.7.2 | Requirement | Sampling tubes that indicate deterioration in performance, such as low recovery rate or nonzero blanks, can be regenerated in the same manner. Check regenerated tubes before re-use as described in 12.4.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics