## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASB 061-2021 1st Ed
ANSI/ASB 061-2021 First Edition
Firearms and Toolmarks 3D Measurement Systems and Measurement Quality Control
Discipline: | Firearms & Toolmarks
Checklist version: | V1.1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Requirements | 4 | Section Title
Developmental Validation (Mandatory) | 4.1 | Requirement | As per ANSI/ASB Standard 063, Implementation of 3D Technologies in Forensic Firearm and Toolmark Comparison Laboratories, a developmental validation shall be completed by at least one organization with appropriate knowledge and/or expertise. The developmental validation of imaging hardware typically consists of identifying and citing previously published scientific literature establishing the underlying imaging technology. The methods defined above of coherence scanning interferometry, confocal microscopy, confocal chromatic microscopy, focus variation microscopy, phase-shifting interferometric microscopy, photometric stereo, and structured light projection all have ample published scientific literature which can be cited to establish the underlying imaging technology.
Deployment Validation (Mandatory) | 4.2 | Section Title
General | 4.2.1 | Requirement | As per ANSI/ASB Standard 063, Implementation of 3D Technologies in Forensic Firearm and Toolmark Comparison Laboratories, the laboratory shall complete and document a deployment validation prior to use in casework. The laboratory shall retain the deployment validation documentation.
Personnel | 4.2.2 | Requirement | Personnel responsible for deployment validation shall at a minimum have a bachelor's degree or equivalent degree with a natural science-based or applied science-based major field of study. To the extent possible, a single individual should conduct all components of the deployment validation to minimize sources of operator uncertainty.
Environmental Conditions | 4.2.3 | Recommendation | To help ensure quality measurements, the instrument should be in a low noise and low vibration environment. Air vents and high traffic areas of the building should be avoided. Some instruments (e.g., interferometry) are extremely sensitive to vibrations and should reside on a vibration isolation table. To determine the “noise floor” of the instrument, see 4.2.5.2.
Instrument Calibrations | 4.2.4 | Requirement | The laboratory shall have the instrument calibrated by the manufacturer or instrument provider upon installation.
Minimum Requirements for Establishing/Verifying Instrument Performance | 4.2.5 | Section Title
General | 4.2.5.1 | Requirement | The deployment validation for a new 3D measurement instrument shall verify and document the instrument’s performance. These tests are required after initial installation and after any subsequent hardware modification, or change that affects the scanning process (e.g., after the installation of a new objective or relocation to a new environment). The deployment validation includes the following.
Instrument Noise Floor Testing | 4.2.5.2 | Requirement | Noise floor testing shall be conducted using a flat surface (e.g., mirror) with flatness better than λ/10, where λ is the wavelength of the light source. Ten measurements shall be performed consecutively without changing the measurement setup. Ten measurements shall also be performed daily over ten days. The calculated average roughness (Sa) value shall be within manufacturer specifications. For 3D instruments that rely on a minimum level of roughness to resolve the surface (i.e., focus variation), the noise floor should be tested using manufacturer suggested standards and protocols.
Instrument Repeatability Measurement | 4.2.5.3 | Section Title
General | 4.2.5.3.1 | Requirement | Using calibrated geometric standards (e.g., sine wave, pitch, step heights), measurements shall be conducted to check the X and Y lateral scales as well as the vertical Z scale. Ten measurements shall be performed consecutively without changing the measurement setup (i.e., not taking the sample off the instrument, not changing the operator, and not varying any measurement parameters). The measurement uncertainty of the repeatability measurements shall overlap with the certified value and uncertainty of the geometric standard used.
X and Y Lateral Scale Check | 4.2.5.3.2 | Requirement | The X and Y lateral scales shall be checked using a geometric standard calibrated for a length scale parameter such as pitch or RSm (wavelength) in standard units.
Z Vertical Scale Check | 4.2.5.3.3 | Requirement | The Z scale shall be checked using a geometric standard calibrated for height in standard units.
Instrument Reproducibility Measurement | 4.2.5.4 | Section Title
General | 4.2.5.4.1 | Requirement | Using calibrated geometric standards (e.g., sine wave, pitch, step heights), check measurements shall be conducted to test the X and Y lateral scales as well as the vertical Z scale. Daily measurements shall be performed over ten days by the same operator. The measurement setup shall be varied each day (i.e., taking the sample on and off the instrument and manually setting up each measurement). The measurement uncertainty of the reproducibility measurements shall overlap with the certified value and uncertainty of the geometric standard used.
X and Y Lateral Scale Check | 4.2.5.4.2 | Requirement | The X and Y lateral scales shall be checked using a geometric standard calibrated for a length scale parameter such as pitch or RSm (wavelength) in standard units.
Z Vertical Scale Check | 4.2.5.4.3 | Requirement | The Z scale shall be checked using a geometric standard calibrated for height in standard units.
Instrument Measurement Uncertainty | 4.2.5.5 | Requirement | The laboratory shall document the instrument’s X, Y, and Z measurement uncertainties in accordance with the “Evaluation of measurement data — Guide to the expression of uncertainty in measurement”. Laboratories may also have to comply with their accrediting body requirements for uncertainty in measurement. These uncertainties can often be determined using the procedures stated in 4.2.5.3 and 4.2.5.4.
Ongoing Performance Checks (Mandatory) | 4.3 | Section Title
General | 4.3.1 | Requirement | As per ANSI/ASB Standard 063, Implementation of 3D Technologies in Forensic Firearm and Toolmark Comparison Laboratories, the laboratory shall document and demonstrate measurement Quality Control (QC) procedures and measurement traceability. The laboratory shall retain the performance check documentation.
Check and Re-check Measurement | 4.3.2 | Requirement | The check and re-check measurements bracket regular measurements. These measurements shall be conducted at the beginning and end of the data acquisition session. These should include at a minimum a check and re-check for the X, Y, and Z scale. The purpose is to ensure that regular measurements are traceable and that the instrument did not drift during the measurements.
Control Chart | 4.3.3 | Requirement | The laboratory shall maintain a control chart tracking all measured re-check measurement values. The control chart shall display the calibrated uncertainty values for each standard and the dated re-check measurement results. Out of spec re-check measurements shall be documented in the control chart.
Quality Control Failure Protocol (Mandatory) | 4.3.4 | Requirement | If the measured check and re-check value falls out of the control limits, a minimum of two repeat measurements shall be made to eliminate the chance of an outlier. The failed check values shall be noted on the control chart. If the error persists, then the hardware shall be removed from service until it can be remedied and a performance validation has been completed. The data collected while in noncompliance of the quality control check (i.e., the compromised measurement session) cannot be guaranteed and shall not be used. The qualified personnel that conducted the instrument validation should, to the best of their ability, diagnose and correct the source(s) of error (e.g., environment, algorithm, protocol, measurement/sample quality, user). If the error diagnosis is inconclusive, the instrument manufacturer or qualified expert should be involved in the investigation.
Measurement Traceability (Mandatory) | 4.4 | Requirement | The geometric standards used for instrument performance validation (see 4.2.5) and quality control measurements (Section 4.3) shall be metrologically traceable to the SI unit of length. Metrological traceability is defined as a result that can be related to a reference through an unbroken chain of calibrations, each contributing to the measurement uncertainty. For more information on traceability, see BIPM, International vocabulary of metrology—Basic and general concepts and associated terms (VIM), JCGM 100:2008. The laboratory shall document the traceability chain of the geometric standards used. This can be supplied by the manufacturer or the accredited calibration laboratory.
Data Exchange Format XML 3D Profile (X3P) (Mandatory) | 4.5 | Requirement | Any imaging hardware that collects data of dimensions greater than two shall have the capability to read and write full resolution raw data in XML 3D Profile (X3P) format in accordance to ISO 25178-72. More information as well as minimum metadata requirements can be found in ANSI/ASB Standard 062, Standard for Topography Comparison Software for Firearm and Toolmark Analysis standard.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics