## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E1588-20
ANSI/ASTM E1588-20
Standard Practice for Gunshot Residue Analysis by Scanning Electron Microscopy/Energy Dispersive X-Ray Spectrometry
Discipline: | Gunshot Residue Analysis
Checklist version: | V1.1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Sample Preparation | 6 | Section Title
Sample Preparation | 6.1 | Recommendation | Once the evidence seal is broken, care should be taken so that no object touches the surface of the adhesive SEM/EDS sample collection stub and that the stub is not left uncovered any longer than is reasonable for transfer, mounting, or labeling.
Sample Preparation | 6.2 | Requirement | The sample collection stub shall be labeled in such a manner that it is distinguishable from other sample collection stubs without compromising the sample; for example, label the bottom or side of the stub.
Sample Preparation | 6.3 | Informational Text | If a non-conductive adhesive was used in the sample collection stub, the sample will need to be coated to increase its electrical conductivity, unless an environmental SEM or variable-pressure/low-vacuum SEM is used for the analysis. Carbon is a common choice of coating material, since it will not interfere with X-ray lines of interest. For high-vacuum SEM, coat the sample sufficiently to eliminate charging of the sample.
Sample Preparation | 6.4 | Requirement | Observe the appropriate procedures for handling and documentation of all submitted samples, for example Practice E1492.
Sample Area | 7 | Section Title
Sample Area | 7.1 | Informational Text | Sample collection stubs for SEMs typically come in one of two diameters: 12.7 mm or 25.4 mm, which yield surface areas of 126.7 mm2 and 506.7 mm2 respectively.
Sample Area | 7.2 | Informational Text | Automated SEM/EDS analysis can enable data collection from nearly the entire surface area of the sample collection stub. Due to the disparity between the shape of the sample collection stub (round) and the SEM field of view search area (square or rectangular), analysis of 100 % of the sample collection area is not always possible in some systems.
Sample Area | 7.2.1 | Recommendation | Analysis of the maximum allowable surface area of the sample is recommended, however, many automated systems can be programmed to terminate the analysis of a stub or series of stubs once a pre-established number of particles have been detected. The decision as to how many particles satisfy the requirements of a particular case should be set out in the laboratory’s standard operating procedures.
Instrument Requirements and Operations | 8 | Section Title
General | 8.1 | Section Title
General | 8.1.1 | Informational Text | Most commercial-grade SEM/EDS systems should be adequate for GSR analysis given that the criteria set forth in 8.2 and 8.3 are met.
General | 8.1.2 | Requirement | Automated data collection of GSR involves some portion of the data collection being controlled by instrument automation software. The extent to which the SEM and EDS systems communicate and are integrated varies according to the manufacturers and the capabilities of the hardware/software. The system shall have the ability to recall stage locations of particles for verification and software for particle recognition.
Scanning Electron Microscope (SEM) | 8.2 | Section Title
Scanning Electron Microscope (SEM) | 8.2.1 | Requirement | The SEM, operating in the backscattered electron imaging mode, shall be configured to detect particles down to at least 1.0 µm in diameter.
Scanning Electron Microscope (SEM) | 8.2.2 | Requirement | The SEM shall be capable of an accelerating voltage of at least 20 kV.
Scanning Electron Microscope (SEM) | 8.2.3 | Requirement | SEM systems shall include: a motorized stage with automated stage control.
Energy Dispersive X-Ray Spectrometry (EDS) | 8.3 | Section Title
Energy Dispersive X-Ray Spectrometry (EDS) | 8.3.1 | Requirement | The detector shall be configured to produce a resolution of better (less) than 150 eV during analysis, measured or extrapolated as the full width at half the maximum height of the Mn Ka peak.
Energy Dispersive X-Ray Spectrometry (EDS) | 8.3.2 | Requirement | At a minimum, the EDS spectrum shall be acquired at 20 eV per channel.
Energy Dispersive X-Ray Spectrometry (EDS) | 8.3.3 | Requirement | Display of the EDS output shall encompass the X-ray lines of analytical utility, with a minimum range of 0–15 keV.
Energy Dispersive X-Ray Spectrometry (EDS) | 8.3.4 | Informational Text | Automated systems will also include software capable of acquiring X-ray spectra for a specified collection time or total X-ray counts.
Energy Dispersive X-Ray Spectrometry (EDS) | 8.3.5 | Requirement | The instrument shall be capable of recording spectra obtained from the analysis of candidate particles. At a minimum, an automated system shall be capable of storing all of the particle location coordinates.
Sample Placement | 8.4 | Section Title
Sample Placement | 8.4.1 | Requirement | Record the positions of the stubs (sample and standard/reference stubs) on the SEM stage when the samples are inserted.
Sample Placement | 8.4.2 | Informational Text | If it is anticipated or required that additional analyses will be needed, it is desirable that the stub can be returned to the same orientation as before its removal. This could consist of marking the side of each stub and aligning it with marks on the microscope stage or by having stubs that fit into the stage in only one position (for example, stubs with a pin that is a half-circle in cross section).
Detection and Calibration | 8.5 | Section Title
Detection and Calibration | 8.5.1 | Informational Text | Particles of GSR are detected by their backscattered electron signal intensity. The absolute signal intensity that a particle produces is related to the electron beam current, mean atomic number, and size of the particle (for particle sizes on the order of the beam diameter). Particles whose mean atomic numbers are high will appear brighter than those of lower mean atomic number composition. As the beam current increases, the amount of signal each particle produces also increases.
Detection and Calibration | 8.5.2 | Recommendation | The brightness and contrast settings (low and high thresholds) of the backscattered electron detector system determine the limits of detection and discrimination of particles based on their mean atomic number. Threshold settings for the backscattered electron signal should be done with a suitable reference sample of known origin (often supplied by the EDS manufacturer) or elemental standards at the same instrumental settings used for sample analyses. This reference sample should, if possible, be in the microscope chamber at the same time as the samples to be analyzed.
Detection and Calibration | 8.5.3 | Recommendation | The backscattered electron detector’s brightness and contrast should be set to include the high atomic number particles of interest and exclude low atomic number particles that are not of interest. Typically, high contrast and low brightness settings provide an adequate range between threshold limits for ease of detection. If the beam current is changed or drifts, the brightness and contrast threshold limits, which were based on the previous beam current, could no longer be compatible with the new conditions and should be readjusted. Beam current measurements can be taken with (but are not limited to) a Faraday cup, a specimen current meter, or monitored by comparing the integrated counts within the same peak in sequentially collected spectra from a known standard.
Quality Control | 8.6 | Section Title
Quality Control | 8.6.1 | Requirement | When conducting automated analysis of GSR, special measures have to be chosen in order to meet common quality management demands. The use of control charts or tracking of instrument parameters and quality control metrics, or both, is recommended. Therefore, as minimum conditions:
Quality Control | 8.6.1.1 | Requirement | Establish a protocol to confirm appropriate instrument operation parameters prior to sample analysis.
Quality Control | 8.6.1.2 | Requirement | Monitor the EDS X-ray energy calibration and SEM beam current stability at least once per batch of samples using appropriate reference materials of known elemental composition such as copper, cobalt, etc.
Quality Control | 8.6.1.3 | Requirement | Analyze a reference material with particles of known size range and composition to test the accuracy of particle detection and identification. It is recommended that the reference material has been prepared and mounted in a manner comparable to the collection method in use by the submitting agency. The reference material can be a sample of GSR from a known source. Additional environmental particles may be added to ensure the inclusion or exclusion of particular classes of particles. Alternatively, a synthetic, GSR standard containing particles comprised of the primary elements of interest may be used for this purpose. The frequency of analysis of this sample shall be subject to guidelines set out in the laboratory’s standard operating procedures.
Quality Control | 8.6.1.4 | Recommendation | The incorporation of environmental or control samples into the analytical protocol is recommended in order to monitor the cleanliness of the sample collection or analytical system, or both. An environmental sample can be prepared in a number of ways: for example, it may be an unused stub that has been prepared contemporaneously with the questioned samples or a sample taken from the sample collection or analytical environments (exposed to the air or as a direct sample collection from clean workspace), or both.
Quality Control | 8.6.1.5 | Recommendation | Care should be taken when transferring samples to and from the instrument to prevent cross contamination. Tools should be cleaned using a solvent and lint free tissue between samples from different sources.
Data Analysis | 9 | Section Title
Definition and Classification | 9.1 | Section Title
Morphology | 9.1.1 | Section Title
Morphology | 9.1.1.1 | Recommendation | Particles identified as characteristic of or consistent with GSR using this method are often spheroid particles, typically between 0.5 μm and 5.0 μm in diameter; the remainder are irregular in shape or vary from 0.5 μm to 100+ μm in size, or both. In general, it is not consistent with the mechanisms of GSR formation to find particles by SEM-displaying crystalline morphology. However, such particles have occasionally been observed in known primer GSR residues. Since morphology can vary greatly, it should never be considered as the only criterion for identification of GSR.
Elemental Composition | 9.1.2 | Section Title
Elemental Composition | 9.1.2.1 | Informational Text | The elemental composition is the most diagnostic property to determine if a particle could be GSR. In some instances, it could provide additional forensic value to compare the elemental composition of the recovered particulate with case-specific known source items, such as the recovered firearm, cartridge cases, or victim-related items.
Elemental Composition | 9.1.2.2 | Recommendation | Particles with unusual elemental compositions could be encountered in case work. In this circumstance, the elemental compositions of these particles should be compared to case-specific sources, such as cartridges or ammunition/firearm test fire deposits.
Definition and Classification | 9.1.3 | Informational Text | Particles classified as characteristic of GSR will have the following elemental composition:
Definition and Classification | 9.1.3.1 | Informational Text | Lead, antimony, barium.
Definition and Classification | 9.1.4 | Informational Text | Particles classified as consistent with GSR will have one of the following elemental compositions:
Definition and Classification | 9.1.4.1 | Informational Text | Lead, barium, calcium, silicon.
Definition and Classification | 9.1.4.2 | Informational Text | Barium, calcium, silicon.
Definition and Classification | 9.1.4.3 | Informational Text | Antimony, barium.
Definition and Classification | 9.1.4.4 | Informational Text | Lead, antimony.
Definition and Classification | 9.1.4.5 | Informational Text | Barium, aluminum.
Definition and Classification | 9.1.4.6 | Informational Text | Lead, barium.
Definition and Classification | 9.1.5 | Informational Text | Particles classified as commonly associated with GSR will have one of the following elemental compositions:
Definition and Classification | 9.1.5.1 | Informational Text | Lead.
Definition and Classification | 9.1.5.2 | Informational Text | Antimony.
Definition and Classification | 9.1.5.3 | Informational Text | Barium.
Definition and Classification | 9.1.6 | Informational Text | The following compositions have been observed from different kinds of ammunition with “lead-free/non-toxic” primers.
Definition and Classification | 9.1.6.1 | Informational Text | Particles that have a composition characteristic of GSR, will have one of the following elemental compositions: (1) Gadolinium, titanium, zinc. (2) Gallium, copper, tin.
Definition and Classification | 9.1.6.2 | Informational Text | Particles with compositions consistent with GSR from different kinds of “lead-free or non-toxic” ammunitions will have one of the following elemental compositions: (1) Titanium, zinc. (2) Strontium.
Definition and Classification | 9.1.7 | Recommendation | It is common for additional elements to become incorporated into particles of these classes. Such particles can contain but not be limited to one or more of the elements: aluminum, silicon, phosphorus, sulfur, chlorine, potassium, calcium, iron, nickel, copper, zinc, zirconium, and tin. The presence of sulfur in classified particles that contain barium without lead should be viewed with caution as barium sulfate is used in many applications unrelated to the discharge of a firearm. Although rare, the presence of iron should be viewed with caution as these particles can be associated with brake dust. The presence of zirconium should be viewed with caution as these particles can be associated with vehicle air bags.
Definition and Classification | 9.1.8 | Requirement | Additional classifications can be developed for specific types of primer compositions not included in the previous sections. Any new classification should aid in differentiating environmentally or occupationally produced particles that could be found in a sample from GSR. An assessment of the significance of these classifications shall be made in consideration of appropriate research and documentation.
X-Ray Analysis – Data Collection and Confirmation | 9.2 | Section Title
X-Ray Analysis – Data Collection and Confirmation | 9.2.1 | Recommendation | An automated SEM/EDS system should collect elemental data from particles whose backscattered electron signal brightness exceeds the desired threshold setting. It should be able to provide hard copy output and electronic storage of the data including, at minimum, stage X and Y coordinates, field of analysis X and Y coordinates, total number of particles detected, and total number of particles classified as GSR.
X-Ray Analysis – Data Collection and Confirmation | 9.2.2 | Requirement | The candidate particles meeting the above criteria (9.1) shall only be considered confirmed after manual relocation of the particles and re-acquisition of the X-ray spectra. The operator shall collect an EDS spectrum from candidate particles by placing the electron beam in spot mode near the center of the particle or raster an area completely within the particle’s volume. X-ray counts should be collected to permit robust manual identification of all elements of interest. If a brief spectral acquisition indicates that the elements are not characteristic of or consistent with GSR acquisition may be stopped. The number of candidate particles to be confirmed shall be subject to guidelines set out in the laboratory’s standard operating procedures.
Documentation | 10 | Section Title
Documentation | 10.1 | Requirement | The following documentation is required for each particle confirmed in 9.2.2:
Documentation | 10.1.1 | Requirement | Images of the particles showing their morphologies (as defined in 9.1.1).
Documentation | 10.1.2 | Requirement | X-ray spectra of the particles, with all relevant elements present clearly identified and labeled.
Documentation | 10.1.3 | Recommendation | The operator/analyst should follow other intra-laboratory protocols for documentation as appropriate.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics