## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASB 017-2018 1st Ed
ANSI/ASB 017-2018 First Edition
Standard Practices for Measurement Traceability in Forensic Toxicology
Discipline: | Toxicology
Checklist version: | V1.2
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
General Measurement Traceability | 4 | Section Title
Background | 4.1 | Section Title
Background | 4.1.1 | Informational Text | Some measurement processes require only a single step, whereas others may require more than one step. Measurement: process of experimentally obtaining one or more quantity values that can reasonably be attributed to a quantity. Importantly, “process” is the first word in the definition of measurement. Measurement relates to the whole process of obtaining a quantity value. Related terms: Quantity: property of a phenomenon, body, or substance, where the property has a magnitude that can be expressed as a number and a reference. Quantity Value: number and reference together expressing magnitude of a quantity. Quantity Value example (Toxicology): Blood alcohol content: 0.158 grams of ethanol per 100 mL of blood.
Background | 4.1.2 | Informational Text | Measurements are made in different ways in testing and calibration processes. A measurement may be: a. a reported test result; b. a reported calibration result; or c. a quantitative decision point (cutoff) that results in a qualitative test report In all of these scenarios, the measurement(s) may be used as the basis for a conclusion, an interpretation, or an opinion. NOTE A measurement may also be made during the testing or calibration process, but not reported to the customer.
Background | 4.1.3 | Informational Text | Measurement traceability can be characterized by the following essential elements. a. Unbroken Chain of Comparisons – A documented system of comparisons with each step having the essential elements of metrological traceability going back to a stated reference acceptable to the parties, usually a national or international standard. b. Documented Measurement Uncertainty – The measurement uncertainty for each step in the traceability chain must be calculated according to defined methods and must be stated so that an overall uncertainty for the whole chain may be calculated. c. Documented Measurement Procedure – Each step in the chain must be performed according to documented, generally-accepted procedures and the results must be documented. d. Technical Competence – The laboratories or bodies performing one or more steps in the chain must maintain and supply evidence of technical competence (e.g., by maintaining appropriate training records, participating in inter-laboratory comparisons, and by demonstrating that they are accredited by a recognized accreditation body). e. Realization of SI Units – The chain of comparisons must, where possible, end at the realization of the International System of Units (SI). f. Documented Calibration Intervals – Calibrations must be repeated at established and appropriate intervals to preserve metrological traceability; and g. Measurement Assurance – A proper measurement assurance program [however named] must be established to ensure the validity of the measurement process and to ensure the calibration status of equipment, reference standards and reference materials.
Requirements for Measurement Traceability | 4.2 | Section Title
General | 4.2.1 | Requirement | Forensic science service providers establish traceability of a measurement process by making one or more measurements using equipment that has been calibrated with established metrological traceability and/or through the use of certified reference materials in the test or calibration method. For equipment and certified reference materials used to establish and maintain measurement traceability, proper handling and storage procedures which meet or exceed manufacturer’s recommendations must be followed.
Calibration Service Provider | 4.2.2 | Requirement | If traceability of a measurement will be established through the calibration of equipment used to make the measurement, then this calibration shall be performed by an appropriately accredited calibration service supplier that, if available, is either: a. a National Metrology Institute (NMI) that is a signatory to the BIPM - CIPM Mutual Recognition Arrangement, with the calibration to be performed listed in Appendix C of the BIPM Key Comparison Database (KCDB); or b. a service supplier accredited to ISO/IEC 17025:2005 by an accrediting body that is a signatory to the International Laboratory Accreditation Cooperation (ILAC) Mutual Recognition Arrangement, with the calibration to be performed listed in a scope of accreditation. If an appropriately accredited calibration service supplier is not available (as defined above), then the forensic science service provider shall perform their own evaluation. This evaluation shall ensure the external service supplier meets the competence, measurement traceability, and measurement capability requirements of ISO/IEC 17025:2005 or in the ILAC Policy on the Traceability of Measurement Results. The forensic science service provider shall keep objective evidence of this evaluation. The forensic science service provider shall perform a re-evaluation of services at least every two years.
Calibration Program | 4.2.3 | Requirement | The forensic science service provider shall identify all equipment in a calibration program, however named, that requires calibration to establish measurement traceability, either totally or partially. The documentation shall: a. establish the interval for calibration; b. delineate if intermediate checks of the calibration status are required between calibrations and if so, the schedule and procedure for those intermediate checks; and c. require calibration of new equipment of this type (e.g., balance, pipette) prior to use of the equipment in testing, calibration or inspection work. See Section 6 for minimum calibration requirements.
Certified Reference Materials | 4.2.4 | Requirement | If traceability of a measurement will be established through reference material, the forensic science service provider shall establish it through the use of one or more certified reference material(s) (CRM). If available, the CRM shall be obtained by the forensic science service provider from a supplier that is either: a. a National Metrology Institute (NMI) that is a signatory to the BIPM - CIPM Mutual Recognition Arrangement, with the CRM to be purchased included in the BIPM key comparison database (KCDB); or b. an accredited Reference Material Producer that is accredited to ISO Guide 34:2009 or ISO 17034:2016 by an accrediting body that is a signatory to a mutual or multilateral recognition arrangement in an ILAC recognized regional accreditation cooperation or the ILAC Mutual Recognition Arrangement, with a scope of accreditation covering the CRM. If a CRM is not available to the forensic science service provider from a supplier that meets the above requirement, then the forensic science service provider shall perform an evaluation of the CRM supplier. This evaluation shall ensure the CRM supplier meets the competence, measurement traceability, and measurement capability requirements of ISO/IEC 17025:2005 or in the ILAC Policy on the Traceability of Measurement Results. The forensic science service provider shall keep objective evidence of this evaluation. The forensic science service provider shall perform a re-evaluation of services at least every two years.
Modifications to Reference Materials | 4.2.5 | Requirement | If a reference material used to establish traceability, whether certified or not, is diluted, such as a stock or working solution, then the equipment used shall be calibrated as delineated in Section 6.
Measurement Traceability Requirements for Toxicology | 5 | Section Title
Measurement Traceability Requirements for Toxicology | 5.1 | Requirement | Forensic toxicology methods are typically categorized as screening, qualitative confirmation/identification, or quantitative. Measurement traceability shall be established by the forensic toxicology laboratory through the use of one or more metrologically-prepared calibrators.10 Additionally, a laboratory may choose to use controls that are metrologically-prepared. 10 Calibrators may be undiluted CRM(s) or calibrators prepared using appropriately calibrated equipment and CRMs, or reference material(s) (RM) shown to be fit for purpose, as defined in 4.2.4.
Measurement Traceability Requirements for Toxicology | 5.2 | Requirement | The following (5.2.1 and 5.2.2) measurement traceability requirements shall be followed by forensic toxicology laboratories for their analytical methods.
Measurement Traceability Requirements for Toxicology | 5.2.1 | Requirement | Measurement traceability shall be established during method validation and routine analysis of cases for: a. all screening procedures with an established decision point concentration; NOTE See 5.2.2 for immunoassay-based screening procedures b. qualitative confirmation/identification methods with an established decision point concentration; c. quantitative methods.
Measurement Traceability Requirements for Toxicology | 5.2.2 | Requirement | Measurement traceability for immunoassay-based screening procedures shall be established during method validation or during routine analysis of cases. Manufacturer kits used in immunoassay-based screening procedures may lack measurement traceability; therefore, laboratories may establish measurement traceability through use of CRMs while initially validating the method and confirming traceability through use of in-house prepared controls.
Equipment | 6 | Section Title
General | 6.1 | Requirement | The equipment listed in Section 6.2 is commonly used in forensic toxicology laboratories and shall be calibrated by appropriately accredited calibration service suppliers that meet those requirements in Section 4.2.2. Documentation of the accredited calibration service suppliers shall be maintained by the forensic toxicology laboratory. Equipment shall be calibrated prior to use, and the laboratory shall have a procedure that includes the frequency of calibration and acceptability/tolerance specifications. Calibration documentation shall include the date the calibration is performed, the calibration status and the date the next calibration is due. Equipment shall be calibrated at a point or within a range consistent with typical use. The forensic toxicology laboratory shall evaluate whether intermediate checks of the calibration status are necessary based on, but not limited to, the frequency of use, work volume, occurrence of unexpected shutdown and equipment maintenance. If intermediate checks are performed, the policy and procedure shall include the frequency and specifications for intermediate checks and actions to be taken when the specifications are not met. Intermediate checks shall be carried out using calibrated equipment (e.g., mass reference standards, equipment used to monitor environmental conditions).
Calibration of Analytical Equipment | 6.2 | Section Title
Analytical Balances | 6.2.1 | Requirement | Analytical balances shall be calibrated at least annually by an appropriately accredited calibration service supplier.
Reference Standards | 6.2.2 | Requirement | Reference standards (e.g., calipers, rulers, mass reference standards) shall be calibrated at least once every three years by an appropriately accredited calibration service supplier. Any adjustments of reference standards shall only be conducted by an accredited calibration service supplier, and the reference standards shall be calibrated before and after any adjustment. Reference standards used to verify the accuracy of equipment shall be dedicated for this purpose, unless the forensic toxicology laboratory has demonstrated that their integrity as reference standards are maintained.
Volumetric Glassware | 6.2.3 | Requirement | Class A volumetric glassware shall be used for the preparation of calibrators and shall be calibrated by an accredited calibration service supplier prior to use. Volumetric glassware used in the preparation of calibrators shall be dedicated for this purpose, and shall be maintained and stored as to protect its integrity. After initial calibration, scheduled recalibration shall recur at least once every ten years by an appropriately accredited calibration service supplier.
Pipettes, Diluters, and Syringes | 6.2.4 | Requirement | All pipettes, pipette diluters, automatic diluters, and syringes used for the preparation of calibrator solutions that require measurement traceability or in sample preparation (e.g. sample aliquoting and other steps that affect overall measurement uncertainty) shall be calibrated at least annually by an appropriately accredited calibration service supplier. Autosampler syringes used for sample introduction to analytical instrumentation (e.g., gas chromatograph, liquid chromatograph, or immunoassay) do not require calibration.
Thermometers | 6.2.5 | Requirement | Thermometers used to verify proper storage of certified reference materials shall be calibrated at least every two years by an appropriately accredited calibration service supplier. Other thermometers that do not significantly affect the accuracy and validity of the test result (e.g., thermometers used in water baths and heat blocks) do not require calibration to establish measurement traceability.
Breath Alcohol Calibration Equipment | 6.2.6 | Requirement | Simulator thermometers, multi-meters and barometers, as applicable, shall be calibrated at least every two years by an appropriately accredited calibration service supplier, as defined in 4.2.2. Sections 6.2.1 to 6.2.5 apply to equipment used in preparation of breath alcohol reference materials used in a breath alcohol calibration method.
Other Equipment | 6.3 | Informational Text | General laboratory equipment used during sample preparation (e.g., centrifuges, rotators, shakers, water baths, evaporators, extraction manifolds, and heating blocks)that does not significantly affect the accuracy and validity of the test result, do not require calibration to establish measurement traceability. A forensic toxicology laboratory may choose to use calibration or verification as a maintenance procedure to ensure proper functioning of the equipment.
Conformance | 7 | Section Title
Conformance | 7 | Requirement | Documentation to verify conformance with the above requirements shall be maintained by the forensic toxicology laboratory for a minimum of five years after expiration and shall be made available to auditors upon request.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics