## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E1967-19
ANSI/ASTM E1967-19
Standard Test Method for the Automated Determination of Refractive Index of Glass Samples Using the Oil Immersion Method and a Phase Contrast Microscope
Discipline: | Trace Evidence
Checklist version: | V1.1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Apparatus | 5 | Requirement
Apparatus | 5.1 | Requirement | Microscope A microscope outfitted for phase contrast and an appropriate long working distance objective (nominally 10–40×) is required.
Apparatus | 5.2 | Requirement | Temperature Control A hot stage is required that connects to a control device with a working range of approximately 26°C to 118°C, having a minimum precision of 0.1°C.
Apparatus | 5.3 | Requirement | Imaging A digital video camera is required for the automated measurements and is mounted to an ocular or photography port of the microscope. The output from the camera (image) is used for automated match point determinations.
Apparatus | 5.4 | Requirement | Illumination A narrow band interference filter is required. For Sodium D measurements 589 ± 5 nm with a band pass of 10 nm is appropriate. The intensity of the illumination is adjusted to give the brightest image possible, without overloading the video camera.
Apparatus | 5.5 | Requirement | Immersion Oils Silicone immersion oils having refractive indices within a specific range are required and are calibrated with the necessary standard reference glasses of known ηDt.
Apparatus | 5.6 | Requirement | Standard Reference Glasses A minimum of three reference ηDt are required to construct the calibration curve for each silicone oil used. When they are commercially available, three or more traceable reference glasses shall be used to construct the calibration curve. In addition, a minimum of one reference glass (control) not used to construct the calibration curve is required to verify the calibration.
Apparatus | 5.7 | Requirement | Computer A computer is required that is equipped with appropriate software and hardware to view the glass fragments, measure edge contrast, and record match temperatures.
Procedure | 6 | Section Title
Procedure | 6.1 | Requirement | Select a number of glass fragments from the known glass sample such that the variability in the refractive index throughout the source can be assessed.
Procedure | 6.2 | Recommendation | A minimum of ten refractive index replicate measurements per known source for most non-tempered glass and twenty for most tempered glasses is recommended.
Procedure | 6.3 | Requirement | Collect replicate measurements of the questioned glass to ensure that it is adequately characterized. A minimum of three replicates is recommended for the comparison of the refractive index of glass.
Procedure | 6.4 | Requirement | Arrange the microscope for optimum illumination and phase contrast. To ensure maximum contrast, make sure the annular illumination ring from the condenser is aligned properly with the phase contrast shift plate, which is located within the objective. The rings are imaged on the back focal plane of the objective and can be observed in a number of ways, the most convenient of which is the use of Bertrand lens or a phase centering telescope.
Procedure | 6.5 | Requirement | Calibrate the necessary ηDT oil as indicated in 7.3.
Procedure | 6.6 | Requirement | After using an appropriate cleaning technique, such as a deionized water and alcohol rinse followed by drying, crush a small fragment of the glass to be measured and deposit a small sample on a clean microscope slide. Immerse this sample in the proper calibrated silicone oil and cover with a cover slip.
Procedure | 6.7 | Requirement | Place the covered slide onto the hot stage and focus the image. The phase ring alignment shall be checked each time that a new slide preparation is inserted into the hot stage to ensure that the phase rings are in alignment.
Procedure | 6.8 | Requirement | To minimize analysis time, set the temperature of the hot stage to within several degrees of the estimated match point. To estimate the match point of a glass, vary the temperature of the hot stage either up or down as appropriate until the match point has been reached and passed. The match point is that point at which the contrast is at a minimum, which corresponds to the disappearance of the edge of interest. With computer-controlled units, the match point will be recorded automatically.
Procedure | 6.9 | Requirement | Using the imaging and controller software, analyze localized portions (commonly referred to as “measurement windows”) of the edge of the glass for changes in refractive index. Place each measurement window across a portion of a focused glass edge and the immersion oil. Begin the automated analysis, which will change the hot stage temperature in a controlled manner. Alternatively, record the video image of the heating and cooling cycles before setting the measurement windows. The computer automatically records the match point temperatures for both cooling and heating cycles for each measurement window and calculates the average of the heating and cooling match point temperatures for every measurement window separately.
Procedure | 6.10 | Informational Text | The ηDT of the glass fragment is measured by calculating the ηDT from the calibration curve (ηDT versus match temperature) based on the average match temperature. For computer-controlled units, this calculation is performed and displayed automatically. The ηDT value will represent the ηDT of the sample at the match point temperature. To obtain the ηDT at ambient temperature the value must be corrected using the dn/dt (change of refractive index with temperature) for that glass. Note that the dn/dt of a particular sample is not usually known for casework glass samples.
Procedure | 6.11 | Requirement | Record all pertinent measurements and instrument data in the case records.
Calibration and Standardization | 7 | Section Title
Calibration and Standardization | 7.1 | Requirement | Check the system calibration each day of use prior to the performance of an analysis.
Calibration and Standardization | 7.1.1 | Requirement | A separate reference glass (control) of known refractive index, distinct from that used for the calibration, for example, National Institute of Standards and Technology (NIST), Schott, Locke, shall be used to verify the calibration curve.
Calibration and Standardization | 7.2 | Requirement | Recalibrate the system any time that the control falls outside the acceptable parameters established by the laboratory or analyst for this procedure.
Calibration and Standardization | 7.3 | Requirement | Calibrate the necessary ηDT oil from a set of three oils of approximately 1.50, 1.53, and 1.55, using reference glasses of known ηDT to ± 0.00001. A minimum of two traceable reference glasses of known ηDT are required to construct the calibration curve for each silicone oil. When they are commercially available, three or more traceable reference glasses shall be used to construct the calibration curve. Once calibrated, the ηDT of the oils can be plotted against the match temperatures to produce a calibration curve for each oil. A preprogrammed protocol within the automated system can be used to perform this function.
Calculation and Interpretation of Results | 8 | Section Title
Calculation and Interpretation of Results | 8.1 | Informational Text | Refractive index measurements are used to test the hypothesis of commonality of the glass from a known broken source and the recovered questioned sample. A refractive index test interval is determined from the broken known glass. Measurements of the questioned glass are then compared to this test interval.
Calculation and Interpretation of Results | 8.1.1 | Recommendation | The analyst should be aware that the number of measurements and the distribution of refractive index values in a glass source affect Type I (false exclusion) and Type II (false inclusion) errors.
Calculation and Interpretation of Results | 8.2 | Informational Text | The procedure to conduct a forensic glass comparison using the recommended test criterion is as follows:
Calculation and Interpretation of Results | 8.2.1 | Requirement | Measure multiple glass fragments as in accordance with 6.2 to determine the minimum and maximum range of the known glass sample (test interval).
Calculation and Interpretation of Results | 8.2.2 | Requirement | Calculate the average refractive index of the questioned sample (6.3).
Calculation and Interpretation of Results | 8.2.3 | Informational Text | If the average refractive index of the questioned sample falls within the minimum value to the maximum value range of the known glass sample, the refractive index of the two glass samples is considered indistinguishable.
Calculation and Interpretation of Results | 8.3 | Informational Text | Interpretation of the frequency of occurrence of a particular refractive index may be assessed with relevant databases of known refractive index values that have been demonstrated to be statistically valid.
Precision and Bias | 9 | Section Title
Precision and Bias | 9.1 | Informational Text | The precision of this test method is based on an interlaboratory study conducted in 2017. Thirty-one laboratories tested three glass samples. Every “test result” represents an individual determination, and duplicate test results were reported by each participant. Practice E691 was followed for the design and analysis of the data; the details are given in a research report.
Repeatability Limit (r) | 9.1.1 | Informational Text | The difference between repetitive results obtained by the same operator in a given laboratory applying the same test method with the same apparatus under constant operating conditions on identical test material within short intervals of time would in the long run, in the normal and correct operation of the test method, exceed the following values only in one case in 20.
Repeatability Limit (r) | 9.1.1.1 | Informational Text | Repeatability limit can be interpreted as maximum difference between two results, obtained under repeatability conditions, that is accepted as plausible due to random causes under normal and correct operation of the test method.
Repeatability Limit (r) | 9.1.1.2 | Informational Text | Repeatability limits are listed in Table 1.
Reproducibility Limit (R) | 9.1.2 | Informational Text | The difference between two single and independent results obtained by different operators applying the same test method in different laboratories using different apparatus on identical test material would, in the long run, in the normal and correct operation of the test method, exceed the following values only in one case in 20.
Reproducibility Limit (R) | 9.1.2.1 | Informational Text | Reproducibility limit can be interpreted as maximum difference between two results, obtained under reproducibility conditions, that is accepted as plausible due to random causes under normal and correct operation of the test method.
Reproducibility Limit (R) | 9.1.2.2 | Informational Text | Reproducibility limits are listed in Table 1.
Bias | 9.1.3 | Informational Text | The bias of a measurement process is a consistent or systematic difference between an average of a set of test results and the accepted reference value for the measured property. An accepted reference value is a value that serves an agreed-upon reference for comparison.
Bias | 9.1.3.1 | Informational Text | Bias is listed in Table 1.
Precision and Bias | 9.1.4 | Informational Text | The above terms (repeatability limit, reproducibility limit, and bias) are used as specified in Practice E177.
Precision and Bias | 9.1.5 | Informational Text | Any judgment in accordance with statements 9.1.1 and 9.1.2 would have an approximate 95 % probability of being correct.
Precision and Bias | 9.2 | Informational Text | The precision and bias statement was determined through statistical examination of 178 test results, from 31 laboratories, on 3 materials.
Precision and Bias | 9.3 | Recommendation | To judge the equivalency of two test results, it is recommended to choose the listed material closest in characteristics to the test material.
Precision and Bias | 9.4 | Requirement | Calculate and document the precision on each individual instrument during the validation of the instrument in the laboratory. The precision is independent of both the accuracy of the temperature measurement and the characteristics of the silicone oil.
Precision and Bias | 9.5 | Requirement | Instrument bias may be introduced from a number of different sources including, but not limited to, calibration procedures, different calibration oils, user practices, and between-instrument variation. The bias of the measurements shall be calculated and documented if databases are used for estimation of frequencies of occurrence or if comparisons using data collected from different instruments are conducted. Because the measurement of the sample ηDT is a direct comparison to the standard reference glasses used, correction for bias need not be calculated for known/questioned comparisons measured using the same calibration on the same instrument.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics