## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E2809-22
ANSI/ASTM E2809-22
Standard Guide for Using Scanning Electron Microscopy/Energy Dispersive X-Ray Spectroscopy (SEM/EDS) in Forensic Polymer Examinations
Discipline: | Trace Evidence
Checklist version: | V1.2
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Sample Preparation | 5 | Section Title
Sample Handling | 5.1 | Section Title
Sample Handling | 5.1.1 | Requirement | Practice E1492, the relevant portions of Guides E1610 and E3260, and the SWGMAT Trace Evidence Quality Assurance Guidelines and Trace Evidence Recovery Guidelines are followed for the collection, handling, and tracking of samples and specimens.
Cleanliness | 5.1.2 | Requirement | Keep the work area and tools used for the preparation of samples free of all materials that could be transferred to the sample. Protect samples prepared for SEM/EDS analysis to minimize possible contributions from the laboratory environment.
Labeling | 5.1.3 | Requirement | Label samples adequately throughout the examination process to ensure that identity and orientation (when relevant) are maintained. This is particularly important when comparison samples are mounted in the same preparation.
Preservation | 5.1.4 | Requirement | When possible, maintain a portion of the evidence in its original, unaltered condition to ensure that adequate sample remains for potential future analyses. In the event that a limited sample size predicates the use of an entire sample, retain the prepared samples as evidence.
Initial Examination | 5.2 | Section Title
Initial Examination | 5.2.1 | Requirement | Record the following minimum information in notes or images for the specific polymer sample being analyzed by SEM/EDS. This information is supplemented through an initial examination by stereomicroscopy. If this information has already been recorded as part of the broader analytical scheme for a sample, this information need not be duplicated.
Initial Examination | 5.2.1.1 | Requirement | Record if the submitted sample is a known or questioned item.
Initial Examination | 5.2.1.2 | Requirement | Record if the sample is suspected (or known) to be a particular type of material (for example, paint or tape).
Initial Examination | 5.2.1.3 | Requirement | Describe the type of sample (for example, chip, smear, tape fragment).
Initial Examination | 5.2.1.4 | Requirement | Determine if the polymer sample is a multilayered product. If so, determine and record the layer structure of the sample. Layer structure determined by means of stereomicroscopy could need to be refined after sample preparation and imaging at higher magnification (by light or electron microscopy).
Initial Examination | 5.2.1.5 | Requirement | Note features that could impact the SEM/EDS analysis (for example, surface imperfections or contaminants, inclusions within the layer).
Visual Inspection for Debris | 5.3 | Informational Text | A stereomicroscope is used to evaluate a sample for the presence of extraneous materials (for example, blood, adhesive from a tape lift, contributions from an underlying substrate). The presence of debris and the approach(es) used to mitigate the impact on sample visualization and elemental analysis results, such as the following:
Physical Removal | 5.3.1 | Informational Text | Under a stereomicroscope, larger particles and some residues are gently scraped or manually picked from the surface of a sample with intent to minimize disruption to the underlying sample surface.
Washing | 5.3.2 | Informational Text | A sample can be washed in a solvent (for example, water, alcohol) with sonication to assist as needed. Prior to washing, a small sample fragment can be subjected to the selected solvent to ensure that the sample is not soluble.
Sectioning | 5.3.3 | Informational Text | The removal of cross-sections permits the exposure of internal surfaces which are not subject to the presence of extraneous materials.
Avoidance | 5.3.4 | Informational Text | In some cases, microanalytical procedures provide imaging capabilities and an analytical volume that permit extraneous debris to be physically avoided during analysis.
Layer Preparation | 5.4 | Requirement | Samples for comparison are prepared under the same conditions whenever possible. The preparation approach, details required to reproduce the preparation process, and differences between sample preparations are recorded. Samples are prepared in a manner that permits the resolution and analysis of individual layers. The following preparations represent a selection of the approaches available. Smears represent an exception and are treated in section 5.4.4.
Hand-Cut Thin-Section Preparation | 5.4.1 | Section Title
Hand-Cut Thin-Section Preparation | 5.4.1.1 | Informational Text | With the aid of magnification (for example, stereomicroscope), thin peels (thin slices through an individual layer) or cross-sections (thin sections which include all layers in a sequence of multilayered polymer samples) can be cut freehand using a scalpel or razor.
Hand-Cut Thin-Section Preparation | 5.4.1.2 | Informational Text | Caution—Polymers can contain layers that are too thin to isolate using the thin peel approach.
Hand-Cut Thin-Section Preparation | 5.4.1.3 | Informational Text | These approaches (thin peels and cross-sections) produce thin sections of varying thickness. Use caution when interpreting and comparing data collected from samples cut freehand as layer thickness differences complicate interpretation.
Stair Step Preparation | 5.4.2 | Section Title
Stair Step Preparation | 5.4.2.1 | Requirement | For multilayered polymer samples, freehand thin peels are sequentially removed from the sample using a scalpel or razor with the aid of a stereomicroscope to produce a “stair step” sample in which a planar surface of each layer within a sample is exposed.
Stair Step Preparation | 5.4.2.2 | Requirement | Due to differences in the vertical height of each stair step, position the sample in the SEM such that the exposed “steps” are oriented with a direct line-of-sight to the EDS detector.
Stair Step Preparation | 5.4.2.3 | Informational Text | Caution—Polymers can contain layers that are too thin to isolate or expose using the stair step approach.
Stair Step Preparation | 5.4.2.4 | Informational Text | This approach exposes layers of varying thickness. Use caution when interpreting and comparing data collected from samples prepared by this approach as both layer thickness differences and contributions from lower layers are possible.
Embedded and Microtome-cut Preparations | 5.4.3 | Section Title
Embedded and Microtome-cut Preparations | 5.4.3.1 | Requirement | Consider and establish the broader analysis scheme and sample preservation requirements prior to embedding a sample as recovery of embedded samples is impossible or difficult.
Embedded and Microtome-cut Preparations | 5.4.3.2 | Requirement | Samples can be embedded in a resin (for example, acrylic, epoxy) to produce a sample that is supported for microtomy. To select an appropriate resin, weigh factors that include: resin composition (to minimize contributions to analysis results); viscosity; curing time; wetting (to ensure that the polymer is firmly encased in the resin); and curing process (impact of heat and UV light on the sample). Record the type of resin used.
Embedded and Microtome-cut Preparations | 5.4.3.3 | Requirement | The position of samples within a mold is recorded to ensure that each sample is unambiguously identifiable.
Embedded and Microtome-cut Preparations | 5.4.3.4 | Requirement | To prepare a block for sectioning, ensure that samples are flat, as opposed to tilted, in the mold.
Embedded and Microtome-cut Preparations | 5.4.3.5 | Requirement | The sample block and respective sample layers are oriented relative to the blade in a manner to minimize potential smearing caused by the passage of the blade through the sample.
Embedded and Microtome-cut Preparations | 5.4.3.6 | Requirement | The embedded sample block (containing one or more samples) is cut on a microtome to produce thin sections.
Embedded and Microtome-cut Preparations | 5.4.3.7 | Informational Text | Caution—Polymers can contain layers that are too thin to isolate during analysis due to excited volume impingement into neighboring layers.
Embedded and Microtome-cut Preparations | 5.4.3.8 | Requirement | Microtome-cut sections from the resin block or the sample within the resin block are analyzed. These thin sections can also be used for other instrumental techniques (for example, infrared spectroscopy). Refer to Guide E2937 for further details.
Embedded and Microtome-cut Preparations | 5.4.3.9 | Requirement | The resin of an embedded sample can contribute to the resulting SEM/EDS sample analysis. If not previously characterized, determine the elemental composition of the resin by SEM/EDS.
Embedded and Microtome-cut Preparations | 5.4.3.10 | Requirement | The sample block is retained as evidence following analysis.
Smears | 5.4.4 | Section Title
Smears | 5.4.4.1 | Requirement | Examine smears by stereomicroscopy or polarized light microscopy for evidence of individual layer remnants. To the extent possible, individual layer remnants are recorded and analyzed.
Smears | 5.4.4.2 | Requirement | When present on a substrate, smears are analyzed in situ (that is, on the substrate) or after isolating them from the substrate.
Smears | 5.4.4.3 | Requirement | Analyze the substrate underlying the smear to account for the presence of elements originating from the substrate in the smeared sample. For example, include analysis of the underlying paint when a paint smear is observed on, and potentially commingled with, another paint system. At a minimum, include the top layer and disrupted layers in the substrate analysis.
Mounting a Sample | 5.5 | Section Title
Mounting a Sample | 5.5.1 | Informational Text | Sample substrates (for example, carbon planchet, aluminum stub) provide physical support for samples to be examined in the SEM and a means by which to ground the sample and reduce charging. Polished beryllium, diamond, or pyrolytic graphite substrates are alternatives worth considering when working with small samples (for example, smears) where it is necessary to recover a particle. These alternatives are highly polished surfaces, exhibit a low background, and are reusable.
Mounting a Sample | 5.5.2 | Requirement | While a variety of sample substrate compositions are available, double-sided conductive carbon adhesive tabs, tape, or sheets are attached to the substrate and provide an adhesive surface for mounting and securing forensic polymer samples.
Mounting a Sample | 5.5.3 | Requirement | Samples are mounted on the substrate such that they are flat and adhered to the substrate.
Mounting a Sample | 5.5.4 | Informational Text | To maximize conductivity of the sample to the substrate and thereby reduce or eliminate sample charging, a line of conductive carbon or silver can be applied between the sample surface and substrate or a conductive layer of carbon (that is, carbon coating) can be deposited. Charging can also be reduced or eliminated through the use of variable pressure mode in an SEM.
Mounting a Sample | 5.5.5 | Requirement | Record the type of substrate along with sample preparation, including any coating applied, as well as the position of samples on the substrate. A digital or hand-drawn map can be used to depict the identity and location of each specimen placed on an SEM stub. An indexing mark on the stub can also be included to assist in sample/location orientation during SEM examination.
Mounting a Sample | 5.5.6 | Informational Text | A spectrum of the sample substrate can also be collected and retained.
Instrumental Operating Conditions | 6 | Section Title
Instrumental Operating Conditions | 6.1 | Recommendation | The following are recommended operating parameters that can be altered to optimize conditions for various analytical needs. As the analyst determines specific analytical needs, actual operating conditions can vary.
Instrumental Operating Conditions | 6.1.1 | Requirement | Samples being compared are analyzed under the same operating conditions.
Instrumental Operating Conditions | 6.1.1.1 | Informational Text | SEM Magnification (if images are collected through the SEM software) and EDS Software Magnification (if images are collected through the EDS software) - Magnification standards for SEMs are commercially available. A magnification tolerance of less than 5% is generally achievable. Refer to Practice E766 for additional information on magnification calibration.
EDS Resolution | 6.1.1.2 | Informational Text | In some software packages, this process can be automated. Alternatively, the resolution can be manually checked by measuring the full width of the manganese (Mn) Ká peak at half of the maximum peak height.
EDS Energy Calibration | 6.1.1.3 | Requirement | Check energy calibration at each time constant used. In some software packages, this process can be automated. For manual measurement, the peak centroid energy of a low-and high-energy peak are measured and checked against reference values.
Instrumental Operating Conditions | 6.1.2 | Requirement | Record the values for each parameter, evaluate them against established tolerances, and compare them to historical values obtained from the same SEM/EDS system.
Instrumental Operating Conditions | 6.1.3 | Requirement | Recalibrate the EDS detector if the measured parameters fall outside of the laboratory defined tolerance. Consult the appropriate operation manual for system specific calibration instructions. Do not use the instrument unless each parameter of the performance check falls within its defined tolerance.
Instrument Setup | 6.2 | Section Title
Instrument Setup | 6.2.1 | Informational Text | Refer to the appropriate operating manual for instrument specific setup instructions.
Instrument Setup | 6.2.2 | Requirement | Record conditions that are necessary to reproduce or review the data. Many of these conditions are automatically captured as metadata within images or spectral data files.
Instrument Setup | 6.2.3 | Informational Text | Variation of the operating values is permitted.
Instrument Setup | 6.2.4 | Requirement | Use the same operating conditions when samples are being compared.
SEM Setup | 6.2.5 | Section Title
Vacuum Mode | 6.2.5.1 | Requirement | It is recommended that samples be analyzed under high vacuum conditions. If variable pressure is used, beam spread can result in contributions from outside the pre-defined raster area such as adjacent layers or the sample substrate. Beam spread increases with higher chamber pressure and reduces spatial resolution for acquired images. In variable pressure mode, record the pressure at which an analysis is conducted.
Beam Saturation | 6.2.5.2 | Requirement | For systems with a filament, attain a stable beam current prior to operation.
Voltage | 6.2.5.3 | Recommendation | For EDS analysis of polymers, an electron accelerating voltage should be chosen to sufficiently excite and allow for the detection and confirmation of the presence of high atomic number elements. Caution should be exercised when using higher beam voltages, as they can damage material with less inorganic content (for example, colorless electrical tape adhesive, automotive clearcoat, duct tape backing). Lower accelerating voltages can be used to excite X-ray transitions from lower elemental electron shells, but the analyst will need to be aware of possible peak overlaps such at Ti K-line and Ba L-line.
Beam Current | 6.2.5.4 | Recommendation | The beam current selected when conducting these analyses are typically in the picoamps to nanoamp range; however, the absolute value is not generally measured or specified in SEM/EDS analyses. It is important to achieve a stable beam current (as per instrument instructions). This is typically achieved by letting the beam stabilize at a specific setting for a period of time (established empirically or by manufacturer recommendation). Therefore, no guidance on the absolute scale of this value is provided in this document; however, the beam current is generally set to provide a balance between image resolution and X-ray output. Adjustments to the beam current can be made by varying a setting that may be described as beam current, spot size, or probe current (varies by manufacturer).
Working Distance | 6.2.5.5 | Requirement | The optimal working distance for collecting EDS X-ray spectra is fixed by the SEM/EDS geometry. Consult the installation specifications to determine the optimal working distance for a given SEM/EDS configuration. Conduct analyses at this optimal working distance unless precluded by specific sample constraints (for example, sample size). Note deviations from the optimal working distance.
Condenser Aperture | 6.2.5.6 | Requirement | Select an appropriate condenser aperture size for imaging and EDS analysis. Check the aperture centering at the start of each analysis session. The SEM operating manual provides instrument-specific centering instructions.
Focus and Stigmation | 6.2.5.7 | Requirement | Ensure that the sample is in focus and that the stigmators have been adjusted prior to imaging and analysis.
Magnification | 6.2.5.8 | Informational Text | Magnification can be adjusted to permit imaging and EDS analysis of a sample at different scales. For instance, lower magnifications can be helpful to record the overall sample or a particular layer, while higher magnification can be used to study a specific inclusion.
EDS Setup | 6.2.6 | Section Title
Deadtime | 6.2.6.1 | Requirement | Set the deadtime through adjustment of SEM beam current and EDS time constant to a desired value (refer to EDS instrument specific operating instructions for guidance). Higher dead times are more likely to produce sum peak artifacts in EDS spectra.
Pulse Processor Time | 6.2.6.2 | Requirement | Set the pulse processor time (for example, time constant) to provide a balance between resolution and total analysis time. Use automatically selected time constants with caution due to the possibility that samples being compared could be analyzed under different time constants.
Live Time | 6.2.6.3 | Recommendation | Set a live time that provides sufficient signal-to-noise to meet the analytical needs should be selected. For example, with paint samples, 100 s live count time can provide sufficient signal for a heavily filled paint layer versus 200 s for a clearcoat with minimal inorganic contribution.
Imaging | 6.3 | Section Title
Imaging | 6.3.1 | Informational Text | Different detectors exist for imaging:
Imaging | 6.3.1.1 | Informational Text | Secondary Electron Imaging, which generally provides topographic information about a sample, is typically used in polymer analyses to examine original surfaces (such as surface deposits or surface textures).
Imaging | 6.3.1.2 | Informational Text | Backscatter Electron Imaging, which provides image contrast that is generally proportional to the average atomic number, is typically used in polymer analyses to examine flat/prepared surfaces to determine the number of layers present within a multi-layer sample or to visualize the size and distribution of inclusions (for example, fillers). Components that exhibit cathodoluminescence display contrast that is dis-proportionally high BSE contrast relative to its atomic number.
Imaging | 6.3.2 | Requirement | Position a sample or layer of interest in the field of view and center it under the electron beam. Adjust the magnification and move the sample to select the area to be imaged. Typically, the sample focus and beam stigmation adjustments are made while viewing the sample at a magnification greater than that to be used for analysis or imaging. Select the appropriate scan speed and acquire a secondary electron or backscatter electron image of the sample.
Imaging | 6.3.3 | Requirement | Evaluate the images for charging effects. If charging is detected, mitigate it by using one of the following strategies:
Imaging | 6.3.3.1 | Requirement | Moving to a different area of the sample.
Imaging | 6.3.3.2 | Requirement | Reducing the accelerating voltage. Changes to accelerating voltage will impact the energy range over which elements will be excited, the excitation efficiency of those elements, and the elements which can be detected. The acquired spectrum will reflect changes in detected elements and peak ratios. Therefore, analyze all samples being compared at the same accelerating voltage.
Imaging | 6.3.3.3 | Requirement | Reduce the beam current since the amount of charging is directly proportional to the energy/current of the electron beam. When necessary, use longer EDS acquisition times and scan times to acquire enough counts to meet the analytical needs.
Imaging | 6.3.3.4 | Requirement | For microscopes equipped with a variable pressure mode, an increase in the chamber pressure (typically 30 to 60 Pa) is sufficient for most non-conductive samples. Use the lowest chamber pressure available that neutralizes sample charging. Beam spread, caused by the collisions of the beam electrons with the gas molecules introduced into the chamber, can unpredictably increase the EDS analytical volume and degrade image quality in direct proportion to the amount of chamber pressure selected.
Imaging | 6.3.3.5 | Requirement | When the above options are insufficient, the sample can be removed and the options to maximize conductivity discussed in 5.5.4 can be revisited.
Imaging | 6.3.4 | Recommendation | As many images as needed to provide adequate support for recording the physical characteristics of the sample and observations that are made should be collected. Useful images can include:
Imaging | 6.3.4.1 | Informational Text | An overview image of a sample to record layer structure.
Imaging | 6.3.4.2 | Informational Text | Images of a layer to record layer thickness and the size and distribution of inclusions within the sample and individual sample layers (that is, compositional texture).
EDS Spectrum Collection | 6.4 | Section Title
EDS Spectrum Collection | 6.4.1 | Requirement | Collect a spectrum to obtain elemental information about a particular sample, layer, or inclusion. Determine the specific collection parameters in order to achieve the goals defined in 4.4.
EDS Spectrum Collection | 6.4.2 | Informational Text | The area sampled by an EDS analysis is defined by one of the following approaches:
Spot | 6.4.2.1 | Informational Text | Collect data at a given spot by positioning the electron beam at a fixed position on a sample.
Raster | 6.4.2.2 | Informational Text | Collect a rastered spectrum by defining the area over which the electron beam will be scanned. The EDS software generally permits a variety of shapes (for example, circle, rectangle, freehand area) for the analyst to choose when defining the size and position of the area rastered and a number of raster scan rates.
Spectral Images (that is, elemental maps) | 6.4.2.3 | Informational Text | Spectral datasets, in which each pixel is associated with an EDS spectrum, are collected from an analyst-defined area (for example, rectangle). Spectral imaging can be used to obtain both elemental composition and distribution from a sampled area. The resolution of a spectral image (in pixels), combined with the microscope magnification, defines the size of each pixel in a spectral image. The image size and magnification are adjusted to ensure that the smallest features of interest can be resolved.
EDS Spectrum Collection | 6.4.3 | Requirement | Determine the area in the sample that will be analyzed based on consideration of analysis goals, sample size, the presence of adjacent layers, and sample heterogeneity (as determined by light microscopy or backscatter imaging).
EDS Spectrum Collection | 6.4.4 | Requirement | Use a sufficient live time to provide reasonable counting statistics. This is instrument, sample, and task dependent, but typically requires 100 to 200 s or enough time to acquire 100 000 total counts in the spectrum.
EDS Spectrum Collection | 6.4.5 | Requirement | Collect replicate analyses as necessary to evaluate variation in the sample.
Spectral Data and Interpretation | 7 | Section Title
Peak Detection | 7.1 | Section Title
Peak Detection | 7.1.1 | Informational Text | A peak is generally considered to be detected when the signal is greater than three times the noise (where noise is the square root of the background). According to Goldstein, et al., actual detection limits are material dependent, but in bulk materials is generally accepted to be about 0.1 wt%.
Peak Assignment | 7.2 | Section Title
Peak Assignment | 7.2.1 | Requirement | Assign peaks to an element (or elements) or another spectral feature (for example, escape peak, sum peak, system peak).
Element Assignment | 7.2.1.1 | Requirement | As possible, identify elements in an EDS spectrum using X-ray peak centroid energy values/ markers for K, L, or M α X-rays; the presence of their associated β X-ray peak centroid energies; and the possible presence of L γ X-ray peaks in a spectrum. Typically, EDS software packages can detect peaks above background, locate their centroid energies, compare those energies with established tables of K, L and M X-ray energies and automatically identify those elements by placing element symbols on those peaks. It is the analyst’s responsibility to manually confirm that each automatically identified element is present in the spectrum by inspecting each assigned element’s peak to ensure that the peak centroid is located at its characteristic energy. In addition, determine that the intensities of each element’s characteristic α, β, γ X-ray peaks, or a combination thereof, are in the proper ratios. Only then can the analyst conclude that a particular element is within a sample spectrum.
Element Assignment | 7.2.2 | Informational Text | An element is considered to be present if a distinctive set of lines is produced or when a single peak occurs at an energy which cannot be mistaken for another element or spectral artifact. Identification of an element is not always possible such as when an element is present at low levels, or the lines required for confirmation are overlapped with the lines of other elements.
Overlap Assignment | 7.2.3 | Section Title
Overlap Assignment | 7.2.3.1 | Informational Text | The K, L, or M X-ray peak energies for specific elements can overlap. Common overlapping element X-rays include: Ti Kβ / V Kα V Kβ / Cr Kα Cr Kβ / Mn Kα Mn Kβ / Fe Kα Fe Kβ / Co Kα Pb Mα / S Kα / Mo Lα Ba Lα / Ti Kα K Kβ / Ca Kα Zn Lα / Na Kα P Kα / Zr Lα Al Kα / Br Lα
Overlap Assignment | 7.2.3.2 | Informational Text | The interpretation of overlapping peaks can be aided by spectral simulation and automated spectral deconvolution tools present in most EDS data collection software. The deconvolution of interferences depends on the detector resolution, time constant, excitation voltage, and concentration of overlapping elements. In addition, the collection of additional data under different conditions (for example, the use of higher voltage to examine higher energy X-ray emission lines, the use of slower time constants to obtain better peak resolution) can help to confirm the presence of certain elements or resolve element peak overlaps.
Artifact Peak Assignment | 7.2.4 | Informational Text | Escape peaks, sum peaks and other system peaks arise from the processing of X-ray signals. Escape peaks and sum peaks are common when a large number of X-rays are produced by an element in the sample. Escape peaks are often visible as peaks whose centroid is located 1.74 keV lower than the peak centroid of a highly concentrated element in the sample. Sum peaks are visible as peaks at twice the energy of a highly concentrated element in the sample. Possible approaches to identification of escape, sum, or system peaks include: collection of a spectrum under different conditions (for example, vary the time constant or beam current) or the use of EDS software algorithms to identify or filter out such peaks.
Peak Labeling | 7.2.5 | Requirement | Label each element that is detected. Denote if the identification is confirmed or tentative.
Documentation | 7.2.6 | Requirement | Record each spectrum with assigned peaks to permit a future review of peak assignments. It is recommended that additional views of a spectrum, with expanded axes, are also recorded to illustrate the presence of minor peaks.
Compound Classification | 7.3 | Section Title
Compound Classification | 7.3.1 | Informational Text | The elements detected in a particular analysis can permit an analyst to record observations about the presence of a particular compound. When using EDS data for this purpose, consider the following:
Compound Classification | 7.3.1.1 | Informational Text | EDS is an elemental technique and does not provide information concerning chemical bonds or crystal structure. As such, certain compounds in a sample can be indicated but not specifically identified. For example, the presence of magnesium, silicon, and oxygen in a sample’s X-ray spectrum can suggest the presence of talc; however, that data does not provide a means to identify the presence of talc as opposed to another magnesium silicate compound.
Compound Classification | 7.3.1.2 | Informational Text | The combination of SEM imaging and EDS analyses (or elemental maps) that are focused on a specific inclusion can provide the means to further define the chemistry of an inclusion. For example, the location of an inclusion that is composed of barium, sulfur and oxygen is a strong indication of barium sulfate. This data interpretation is supported by discipline-specific knowledge that pigment-grade barium sulfate is commonly added to polymeric formulations.
Compound Classification | 7.3.2 | Requirement | Where possible, evaluate data from SEM/EDS and orthogonal techniques such as infrared or Raman spectrometry in concert. For example, it is anticipated that a layer that contains rutile titanium dioxide, identified by extended range infrared or Raman spectrometry, would contain titanium and oxygen by SEM/EDS. Refer to Guides E2937 and E3085 for further details.
Compound Classification | 7.3.3 | Requirement | Record the rationale to support the above determinations in the case notes.
SEM/EDS Spectral Comparison | 7.4 | Section Title
SEM/EDS Spectral Comparison | 7.4.1 | Recommendation | SEM/EDS spectral comparisons should be conducted between spectra collected using similar sample preparations, similar sample characteristics (for example, thickness, topography), and similar instrumental parameters, as appropriate.
SEM/EDS Spectral Comparison | 7.4.2 | Informational Text | SEM/EDS spectra are compared and interpreted based on the observation of spectral differences, or lack thereof, between the sets of SEM/EDS data.
SEM/EDS Spectral Comparison | 7.4.2.1 | Informational Text | Spectral overlay can be used for comparing data where the presence or absence of peaks, peak shapes, and relative intensities are all considered in the evaluation as to whether exclusionary differences exist between compared samples.
SEM/EDS Spectral Comparison | 7.4.2.2 | Informational Text | When assessing differences between spectra, consider sample limitations (for example, small samples, thin layers, dirty samples, sample smears that eliminate layer structure) and instrumental limitations (for example, sampling size, limits of detection).
SEM/EDS Spectral Comparison | 7.4.3 | Informational Text | Possible reasons for spectral differences include dissimilar sample characteristics (for example, size, thickness, surface topography), heterogeneity, contribution from extraneous materials, or origination from different source materials. Additional samples can provide supplemental data to assist in assessing such differences.
SEM/EDS Spectral Comparison | 7.4.4 | Informational Text | If suitable SEM/EDS spectra are produced, comparisons can provide information regarding the potential relationship between the sources of the samples.
SEM/EDS Spectral Comparison | 7.4.4.1 | Informational Text | When exclusionary differences are observed between compared spectral features, the sources of the samples are considered distinguishable by SEM/EDS. Exclusionary differences in spectral comparisons: (1) are outside the variability of spectra originating from the same source; and (2) cannot be explained by considerations such as sample heterogeneity, contamination, different sample conditions, or different sample histories.
SEM/EDS Spectral Comparison | 7.4.4.2 | Informational Text | When no exclusionary differences are observed between compared spectral features, the sources of the samples are considered indistinguishable by SEM/EDS. Differences that are not considered exclusionary: (1) are within the variability of spectra originating from the same source; or (2) can be explained by considerations such as sample heterogeneity, contamination, different sample conditions, or different sample histories. If no exclusionary differences are observed in a SEM/EDS comparison, samples can be analyzed by other analytical techniques to provide additional information about the potential relationship between the sources of the samples.
SEM/EDS Spectral Comparison | 7.4.5 | Informational Text | SEM/EDS spectral comparison is one part of a multi-analytical comparative approach. SEM/EDS data alone can be used to distinguish the sources of compared samples but are otherwise not used independent of data obtained from other analytical techniques to reach an overall opinion regarding the potential relationship between the sources of the samples. An overall opinion that sources are indistinguishable is only reported when no exclusionary differences are observed in the analytical techniques that were applied.
Documentation | 8 | Section Title
Documentation | 8.1 | Requirement | Record the details necessary to support the interpretations made from each comparison (Practice E620).
Documentation | 8.2 | Recommendation | Case notes should be sufficient to allow an independent analyst to understand and evaluate all the work performed, independently analyze and interpret the data, and draw conclusions.
Documentation | 8.3 | Informational Text | A description of the evidence analyzed by SEM/EDS, sample preparations, the analytical instrumentation used, and its operating parameters are included in the case notes, case record, or otherwise recorded in accordance with laboratory procedures.
Documentation | 8.4 | Informational Text | Refer to Practices E1492, E620, and ISO 17025 for further guidance.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics