## Instructions for Use
AAFS Standards Resources and Training
This checklist is made possible through the following financial assistance award 70NANB24H078 awarded to AAFS from U.S. Department of Commerce, National Institute of Standards and Technology Each checklist is specific to a version of a standard, which is hyperlinked in the checklist. Not all sections of a standard are included in a checklist. A checklist is a supplementary tool and should not be used independently. To properly use a checklist, reference the complete standard. Go to https://www.nist.gov/organization-scientific-area-committees-forensic-science/access-standards for information on how to access published standards. A checklist does not provide interpretation guidance. If you have questions related to interpretation, please contact the standard developing organization (SDO). The SDO may have standard specific training available.
A checklist can be used for one or both of the following purposes: • to evaluate the implementation of a standard, or • to audit (internal or external) conformance with requirements in the standard and an FSSP’s management system
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions and an on-demand training webinar.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## OSAC Proposed Std 2022-S-0040
OSAC Proposed Standard 2022-S-0040, Version 2.0
Standard for Internal Validation of DNA Extraction Methods
Discipline: | Biology/DNA
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Requirements | 4 | Requirement | ANSI/ASB Standard 038, Standard for Internal Validation of Forensic DNA Analysis Methods, shall be used in conjunction with this document because ANSI/ASB Standard 038 provides the foundational validation requirements upon which additional specific requirements, such as this document, will be based. NOTE 1: The recovered DNA yields, degradation indices, and autosomal/Y-chromosome ratios may be used to assess the performance of the DNA extraction method, which includes the comparison of a new method to a previous one. Statistically significant differences may not be cause for rejection of a new extraction method since the new method could be faster, easier to use, automatable, or the resulting DNA extracts more efficiently amplified. NOTE 2: Refer to Appendix A of the Best Practice Recommendations for a cross-function study table where data from different studies may be used to meet the requirements under another category. NOTE 3: Some of the studies below may not be applicable based on the laboratory’s validation plan or intended purpose for the extraction method. For example, a laboratory may not need to assess multiple biological sample types or substrates.
Requirements | 4.1 | Requirement | The laboratory shall evaluate the efficacy of the DNA extraction method using downstream validated laboratory testing procedures. If previously internally validated procedures are not available, then procedures intended for future use in casework should be used. NOTE: For the assessment of the DNA extraction method using downstream procedures that have not been internally validated, the laboratory may rely on developmental validations, manufacturer recommended settings and methods, or validated procedures from other laboratories.
Requirements | 4.2 | Requirement | The laboratory shall design experiments and define acceptance criteria, where applicable, based, at a minimum, on the requirements listed in 4.3 through 4.8 below.
Requirements | 4.3 | Requirement | The laboratory shall demonstrate concordance by comparison to published developmental validation results or previously developed internal validation data.
Requirements | 4.3.1 | Requirement | DNA typing results of one sample from five different individuals obtained from the new DNA extraction chemistry shall be compared to results obtained from a previously internally or externally validated DNA extraction method.
Requirements | 4.3.2 | Requirement | Discordant typing results observed when comparing identical genetic markers using sample(s) from the same individual shall be documented in the final validation document. NOTE: If possible, a potential explanation for the discordant typing results supported by data or other research studies should be included in the final validation document.
Requirements | 4.4 | Requirement | The laboratory shall evaluate the DNA extraction method using known references, casework-like samples, and substrates representative of those typically analyzed by the testing laboratory.
Requirements | 4.4.1 | Requirement | The laboratory shall assess if the DNA extraction method is fit for purpose by using different biological sample types, substrates, and sample conditions (e.g., inhibited samples, degraded samples). At a minimum, the assessment shall include the requirements stated in 4.4.1.1 through 4.4.1.4.
Requirements | 4.4.1.1 | Requirement | For each identified biological sample type [e.g., blood, semen, saliva, trace (epidermal cells), hairs, bones, tissue], a minimum of one sample from two different individuals shall be tested in triplicate (for a total of six extractions for each biological sample type). Average yield of recovered DNA and standard deviations shall be calculated.
Requirements | 4.4.1.2 | Requirement | For a substrate study, at least one of the biological sample types shall be placed on a variety of different substrates expected to be encountered by the laboratory (minimum of five different substrates) and tested in triplicate. Average yield of recovered DNA and standard deviations shall be calculated.
Requirements | 4.4.1.3 | Requirement | For sample condition studies, the following tests shall be performed: a) At least one of the biological sample types shall contain a known inhibitory substance(s) expected to be encountered by the laboratory [e.g., heme, humic acid (dirt)]. The inhibited sample(s) shall be tested in triplicate, at a minimum, and a qualitative assessment of the presence of inhibition shall be performed. b) At least one of the biological sample types shall be compromised and demonstrated to contain degraded DNA. The biological sample(s) shall be tested in triplicate, at a minimum, and qualitatively assessed.
Requirements | 4.4.1.4 | Requirement | If the results from biological samples fail to meet defined laboratory acceptance criteria, then additional samples shall be tested. The number of additional samples tested depends on the outcome of the initial tests and the evaluation of possible reasons the data were outside the acceptance criteria.
Requirements | 4.4.2 | Requirement | If the laboratory tests samples containing semen, the following tests shall be performed: a) At least one neat semen sample tested in triplicate. b) At least one semen sample mixed with at least two other biological sources (e.g., vaginal swab with semen, blood with semen, saliva with semen) expected to be encountered by the laboratory, tested in triplicate.
Requirements | 4.4.3 | Requirement | The laboratory shall compare the results of known and casework-like samples extracted using the method being validated to the results from previously validated DNA extraction methods, where applicable.
Requirements | 4.5 | Requirement | The laboratory shall determine the susceptibility of the DNA extraction method to cause degradation or to inhibit PCR amplification.
Requirements | 4.5.1 | Requirement | At least two different biological sample types shall be tested in triplicate and the extracts taken through the complete DNA typing process. The resulting DNA profile data shall be qualitatively evaluated for apparent degradation and inhibition as compared to other extraction procedure(s).
Requirements | 4.5.2 | Requirement | The biological samples used shall not contain any known endogenous PCR inhibitors and shall be expected not to be degraded prior to extraction.
Requirements | 4.6 | Requirement | The laboratory shall conduct a sensitivity study to determine the limits of the DNA extraction method for obtaining DNA sufficient for analysis.
Requirements | 4.6.1 | Requirement | A dilution series of at least one biological sample being extracted consisting of three or more dilutions (e.g., neat, 1:10, 1:100), shall be performed and tested in triplicate.
Requirements | 4.6.2 | Requirement | Average yield of recovered DNA and standard deviations shall be calculated.
Requirements | 4.6.3 | Requirement | The samples shall be taken through the complete DNA typing process and the profiles quantitatively assessed.
Requirements | 4.7 | Requirement | The laboratory shall determine the precision (i.e., the reproducibility and repeatability) of the DNA extraction method by calculating the average DNA yields or concentrations and standard deviations. NOTE: Precision studies can use data obtained from multiple studies within the validation.
Requirements | 4.8 | Requirement | The laboratory shall determine the susceptibility of the DNA extraction method to the introduction of exogenous DNA during the extraction process. The samples shall be taken through the complete DNA typing process.
Requirements | 4.8.1 | Requirement | A laboratory shall define what constitutes DNA contamination.
Requirements | 4.8.2 | Requirement | A laboratory shall calculate the DNA contamination rate using reagent blanks and samples containing DNA and use these data to define the expected rate of the method in casework.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed | Clause involves options; this option was not chosen | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Forensic Document Examination
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics