## Instructions for Use
AAFS Standards Resources and Training
This checklist is made possible through the following financial assistance award 70NANB24H078 awarded to AAFS from U.S. Department of Commerce, National Institute of Standards and Technology Each checklist is specific to a version of a standard, which is hyperlinked in the checklist. Not all sections of a standard are included in a checklist. A checklist is a supplementary tool and should not be used independently. To properly use a checklist, reference the complete standard. Go to https://www.nist.gov/organization-scientific-area-committees-forensic-science/access-standards for information on how to access published standards. A checklist does not provide interpretation guidance. If you have questions related to interpretation, please contact the standard developing organization (SDO). The SDO may have standard specific training available.
A checklist can be used for one or both of the following purposes: • to evaluate the implementation of a standard, or • to audit (internal or external) conformance with requirements in the standard and an FSSP’s management system
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions and an on-demand training webinar.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E2881-25
ANSI/ASTM E2881-25
Standard Test Method for Extraction and Derivatization of Fatty Acid-Based Products from Fire Debris and Liquid Samples with Analysis by Gas Chromatography-Mass Spectrometry
Discipline: | Fire Debris & Explosives
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Apparatus | 6 | Section Title
Apparatus | 6.1 | Requirement | Gas Chromatograph— A chromatograph capable of using capillary columns and being interfaced to a mass spectrometer.
Apparatus | 6.1.1 | Requirement | Inlet System— A inlet system that can be operated in either split or splitless mode with capillary columns; the inlet system may use on-column technology.
Apparatus | 6.1.2 | Requirement | Column— A non-polar capillary, bonded phase, methylsilicone or phenylmethylsilicone column or equivalent, or a polar capillary, bonded phase, such as a cyanopropyl-based FAME-specific column, or equivalent, can be used to determine the presence of fatty acids.
Apparatus | 6.1.2.1 | Informational Text | A polar capillary, bonded phase column, such as a cyanopropyl-based FAME-specific column, is used to perform comparative analysis between specimens of questioned neat liquids, debris extracts, or both.
Apparatus | 6.1.2.2 | Informational Text | Any column length or temperature program conditions can be used for the polar column provided that each component of the reference mixture (see 7.8) is adequately separated.
Apparatus | 6.1.3 | Requirement | GC Oven— A column oven capable of reproducible temperature program operation in the range from 50 °C to 300 °C.
Apparatus | 6.2 | Requirement | Mass Spectrometer— A mass spectrometer capable of acquiring mass spectra from m/z 40 to m/z 400 with unit resolution or better, with continuous data output.
Apparatus | 6.2.1 | Requirement | Sensitivity and Resolution— A system capable of detecting each component of the reference or test mixture (see 7.8) and providing sufficient ion intensity data to identify each compound, either by computer search of a mass spectral library or by comparison with reference spectra.
Apparatus | 6.3 | Requirement | Data Station— A computerized data station capable of storing time sequenced mass spectral data files of analyzed specimens.
Apparatus | 6.3.1 | Requirement | Data Handling— A system capable of performing, either through its operating system or by user programming, various data handling functions, including input and storage of data files, generation of extracted ion profiles, searching data files for selected compounds, and qualitative and semi-quantitative compound analysis.
Apparatus | 6.3.2 | Requirement | Mass Spectral Libraries— A system capable of retrieving a specified mass spectrum from a data file and comparing it against a library of mass spectra available to the data system. This capability is considered an aid to the analyst, who will use it in conjunction with chromatographic data and known reference materials to identify unknown compounds.
Apparatus | 6.4 | Requirement | Syringes— A syringe capable of introducing a specimen size in the range from 0.1 μL to 10.0 μL.
Reagents and Materials | 7 | Section Title
Reagents and Materials | 7.1 | Requirement | Purity of Reagents— Reagent grade or better chemicals shall be used in all tests. Unless otherwise indicated, it is intended that all reagents conform to the specifications of the Committee on Analytical Reagents of the American Chemical Society, where such specifications are available. Other grades can be used, provided it is first ascertained that the reagent is of sufficiently high purity to permit its use without lessening the accuracy of the determination.
Reagents and Materials | 7.2 | Requirement | Solvent— A suitable solvent, such as n-pentane, n-hexane, or n-heptane.
Reagents and Materials | 7.2.1 | Informational Text | Solvent purity is evaluated prior to use with specimens. The solvent is analyzed using the same conditions as used for specimens in accordance with the procedure described in this test method.
Reagents and Materials | 7.3 | Requirement | Derivatization Reagent— A suitable derivatization reagent, such as a 2N potassium hydroxide (KOH) in methanol solution (base-catalyzed) is used to convert the triglycerides to FAMEs.
Reagents and Materials | 7.4 | Requirement | Drying Agent— A suitable drying agent, such as anhydrous sodium sulfate.
Reagents and Materials | 7.5 | Requirement | Filter Apparatus— Free of extractable hydrocarbons, oils, and fats.
Reagents and Materials | 7.6 | Requirement | Evaporation Accelerants— Compressed dry nitrogen, filtered air, or inert gas used in the concentration of solvent extracts.
Reagents and Materials | 7.7 | Requirement | Carrier Gas— Helium or hydrogen of purity 99.995 % or higher.
Reagents and Materials | 7.8 | Requirement | Reference Mixture— The reference mixture consists of a minimum of the following FAMEs: C16:0, C18:0, C18:1, C18:2, C18:3. The mixture should contain approximately equal parts by weight of the required FAMEs in the chosen solvent or a traceable commercially available reference mixture. The final solution is prepared by diluting the above mixture such that the concentration of each component is no greater than 0.005 % weight/volume (0.05 mg/mL) in the chosen solvent. A typical chromatogram including components of the reference mixture on a typical non-polar fire debris column is shown in Fig. 1. A typical chromatogram including components of the reference mixture on a FAME-specific column is shown in Fig. 2.
Reagents and Materials | 7.9 | Requirement | Reference Oils and Fats— Oils and fats should be available for comparison and identification purposes.
Reagents and Materials | 7.9.1 | Informational Text | FAMEs derived from reference oils and fats are diluted approximately 1:200 in an appropriate solvent and derivatized using the same procedure that will be used on the debris and liquid samples. Depending on the column capacity and injection technique, derivatized oil and fat solutions can be concentrated to ensure detection of minor compounds.
Reagents and Materials | 7.1 | Requirement | Glassware or Labware— Clean glassware (beakers, test tubes, and vials) or disposable labware free of extractable hydrocarbons, oils, and waxes.
Quality Assurance and Quality Control | 8 | Section Title
Quality Assurance and Quality Control | 8.1 | Informational Text | Refer to Test Method E1618 for equipment calibration and maintenance procedures.
Quality Assurance and Quality Control | 8.2 | Informational Text | Refer to Test Method E1618 for performance criteria for both the polar and non-polar columns.
Quality Assurance and Quality Control | 8.2.1 | Requirement | For non-polar columns, C18:1, C18:2, and C18:3 shall be sufficiently resolved to effect individual identification of each by comparing GC retention time and mass spectral data to a known reference material.
Quality Assurance and Quality Control | 8.2.2 | Requirement | For polar columns, each component of the reference mixture in 7.8 shall have baseline resolution.
Quality Assurance and Quality Control | 8.3 | Informational Text | Refer to Practice E3255 for other quality assurance procedures.
Sample Handling Procedure | 9 | Section Title
Sample Handling Procedure | 9.1 | Requirement | Observe the appropriate procedures for handling and documentation of all submitted samples as described in Practices E1492 and E3255 and Guide E3245.
Sample Handling Procedure | 9.2 | Requirement | Analyze specimens of appropriate dilution on a GC-MS system. Follow sample handling procedures in Test Method E1618.
Sample Handling Procedure | 9.3 | Requirement | Maintain a reference collection that includes known oils and fats that have been prepared and analyzed in the same manner as the specimens.
Sample Handling Procedure | 9.4 | Informational Text | See Annex A1 for sample and extract storage guidance.
Procedure | 10 | Section Title
Procedure | 10.1 | Requirement | If ignitable liquid analysis is needed, perform the ignitable liquid exam first on the entire sample, unless the items are not packaged in vapor-tight containers. Methods for isolating ignitable liquid residues (ILR) from fire debris for analysis are described in Practices E1386, E1388, E1412, E1413, E2154, and E3189.
Procedure | 10.1.1 | Recommendation | Warning—Fire debris samples and extracts that potentially contain triglycerides and free fatty acids should not be subjected to temperatures above 90 °C or evaporated to dryness prior to derivatization, as both of these may change the composition of the extract
Procedure | 10.1.2 | Informational Text | If Practice E1386 is to be used for ILR analysis, an aliquot of the solvent rinse from the FAMEs analysis can be obtained prior to derivatization, concentrated and analyzed using Test Method E1618.
Procedure | 10.2 | Requirement | Following ignitable liquid analysis, perform extraction for oils and fats. The system shall be dry (free of water) and have an excess of methanol present for the reaction to proceed to completion.
Procedure | 10.2.1 | Section Title | Derivatization of Liquid Samples:
Procedure | 10.2.1.1 | Requirement | Dilute liquid in appropriate solvent (see 7.2) using at least 25 μL to 50 μL (1 drop) of liquid to 10 mL of solvent. Use an equal amount of a reference oil as a positive control to process along with the specimen(s).
Procedure | 10.2.1.2 | Requirement | Add at least 0.5 mL of 2N KOH in methanol to the diluted liquid and vigorously mix (vortex mixing or shaken).
Procedure | 10.2.1.3 | Requirement | Allow the layers to separate (centrifuge if necessary), and remove the top (organic) layer. Filter the organic layer as needed with 0.45 μm filters (or finer) designed for organic extracts, such as polytetrafluoroethylene (PTFE) membrane filters.
Procedure | 10.2.1.4 | Requirement | Analyze the derivatized extract by GC-MS using either a non-polar capillary column or a polar capillary, bonded phase column as specified in 6.1.2, using the conditions described in Fig. 1 or Fig. 2. Conditions other than those listed in Figs. 1 and 2 can be used if they have been validated or verified in accordance with Practice E2549 and they provide sufficient sensitivity and resolution to detect and separate the compounds in the test mixture.
Procedure | 10.2.2 | Section Title | Debris Samples:
Procedure | 10.2.2.1 | Requirement | Only extract the portion of the fire debris sample suspected to contain fats or oils. Transfer the selected portion to a clean container (free of extractable hydrocarbons, fats and oils) and add sufficient solvent to thoroughly moisten the fire debris sample. Mix the solvent and debris for several minutes. Simple rinsing of nonporous surfaces does not adequately separate residues. Extracting the entire fire debris sample should be avoided whenever possible as this test is destructive.
Procedure | 10.2.2.2 | Requirement | Remove the extract from the debris. If water is present, after removal of the extract from the debris, allow the solvent and water to separate. Decant the solvent layer into a clean beaker and add a drying agent (for example, anhydrous sodium sulfate) to the sample extract until the drying agent flows freely without clumping. The amount of drying agent added will vary depending on the amount of water present. Before proceeding to the next step, remove the extract from the drying agent and transfer the extract to a new container.
Procedure | 10.2.2.3 | Requirement | Derivatize all or a portion of the extract (typically no more than 10 mL) by either the base-catalyzed or acid-catalyzed method as described above (see 10.2.1.2).
Procedure | 10.2.2.4 | Requirement | Use a reference oil sample (see 10.2.1.1) diluted in the same solvent as was used to extract the debris as a positive control to process along with the specimen(s).
Procedure | 10.2.2.5 | Requirement | After derivatization, allow the layers to separate (centrifuge if necessary), and remove the top (organic) layer. Filter the organic layer as needed with 0.45-μm filters (or finer) designed for organic extracts, such as PTFE membrane filters.
Procedure | 10.2.2.6 | Requirement | Analyze the derivatized extract by GC-MS using either a non-polar capillary column or a polar capillary, bonded phase column as specified in 6.1.2, using the conditions described in Fig. 1 or Fig. 2. Conditions other than those listed in Figs. 1 and 2 can be used if they have been validated or verified in accordance with Practice E2549 and they provide sufficient sensitivity and resolution to detect and separate the compounds in the test mixture.
Procedure | 10.2.3 | Requirement | Concentrate extracts before or after derivatization, as needed, in a chemical fume hood. Evaporate at room temperature to a suitable final volume. Compressed dry nitrogen, filtered air, or an inert gas can be used to accelerate evaporation. Heat is not recommended for concentrating prior to derivatization as the oxidative polymerization of small quantities of extracted fatty acids and triglycerides could be accelerated by heat. Using heat to concentrate after derivatization is acceptable as polymerization does not occur with FAMEs.
Data Analysis and Interpretation | 11 | Section Title
Data Analysis and Interpretation | 11.1 | Requirement | Identify FAMEs in liquids and fire debris extracts by comparing GC retention time and mass spectral data to certified standards, a reference mixture (see 7.8), or an oil or fat reference material. Document the comparison in accordance with Test Method E1618.
Data Analysis and Interpretation | 11.1.1 | Requirement | In order to identify a fatty acid-based product, the FAMEs should be at least one order of magnitude greater than the average abundance of the fire debris sample matrix peaks. Use caution when identifying fatty acids when only saturated fatty acids, such as C16:0 and C18:0, are present. NOTE 1—The presence of FAMEs does not necessarily indicate that a foreign fatty acid-based product is present. These can be inherent to the scene and found in low abundance.
Data Analysis and Interpretation | 11.2 | Section Title | Extracted Ion Profiling (EIP):
Data Analysis and Interpretation | 11.2.1 | Requirement | Use extracted ion profiles to elucidate FAME peaks from other interferences in the chromatograms.
Data Analysis and Interpretation | 11.2.2 | Requirement | Use a data station to extract and draw extracted ion profiles (mass chromatograms) for major ions characteristic of FAMEs.
Data Analysis and Interpretation | 11.2.3 | Informational Text | Typically ions 67, 74, and 79 or the molecular weight ion for the FAME of interest can be summed together to form a useful EIP. See Table 1 for a list of common fatty acids and the molecular weights of the resulting methyl esters.
Data Analysis and Interpretation | 11.2.3.1 | Informational Text | The ions 67 and 79 are indicators for methyl esters of linoleic acid (methyl octadecadienoate) and linolenic acid (methyl octadecatrienoate), which are acids known to be involved with self-heating oils.
Data Analysis and Interpretation | 11.3 | Section Title | Interpretation of FAME Results:
Data Analysis and Interpretation | 11.3.1 | Informational Text | The type of fatty acids recovered can be used to infer the self-heating tendency of the oil recovered from the sample. Quantitation of the fatty acids is not required and is beyond the scope of this test method.
Data Analysis and Interpretation | 11.3.2 | Informational Text | When only saturated fatty acids (for example, C16:0 and C18:0) are present, an oil or fat has no tendency toward self-heating.
Data Analysis and Interpretation | 11.3.3 | Informational Text | When C18:1 and C18:2 are present without C18:3, an oil or fat has a low tendency toward self-heating.
Data Analysis and Interpretation | 11.3.4 | Informational Text | When C18:2 and C18:3 are present and they are not the most abundant fatty acids, an oil or fat has a moderate tendency toward self-heating.
Data Analysis and Interpretation | 11.3.5 | Informational Text | When C18:3 is the most abundant fatty acid, an oil or fat has a high tendency toward self-heating.
Report Wording | 12 | Section Title
Report Wording | 12.1 | Informational Text | Refer to Practice E620 and Test Method E1618, Section 12, for general information on report writing.
Report Wording | 12.2 | Informational Text | Suggested wording for Liquid, Solid, and Debris Samples Containing FAMEs:
Report Wording | 12.2.1 | Informational Text | Fatty acids were identified in Exhibit 1. Fatty acids are naturally occurring and are found in oils and fats. The fatty acids are consistent with known oils or fats; however, the original source could not be determined.
Report Wording | 12.2.1.1 | Informational Text | Oil (or fat) with very slight or low tendency toward spontaneous ignition was identified. This result does not eliminate the possibility that a substance with a higher tendency toward self-heating was present prior to the fire.
Report Wording | 12.2.2 | Informational Text | Exhibit 1 contained fatty acids, which are typical components of oils or fats, with a (no, low, moderate, or high) degree of unsaturation indicating a (no, low, moderate, or high) tendency of self-heating.
Report Wording | 12.2.2.1 | Informational Text | The above characterization of any oils or fats identified is appropriate only in cases in which the fatty acids are known to be unreacted.
Report Wording | 12.2.3 | Informational Text | In the above-mentioned wording, fatty acids can be specifically identified by their chemical name(s) or common name(s).
Report Wording | 12.3 | Section Title | Suggested Wording for Negative Reports:
Report Wording | 12.3.1 | Section Title | No Fatty Acids Identified:
Report Wording | 12.3.1.1 | Informational Text | No fatty acids were identified in the sample, which indicates that no oil or fat was present in the sample. It should be noted that due to the destructive nature of fire, there are no test methods that can exclude the possibility of oils or fats being consumed or polymerized during the fire or that can determine all of the substances that were present prior to the fire.
Report Wording | 12.4 | Section Title | Suggested Wording to Include the Interpretation of the Fatty Acids:
Report Wording | 12.4.1 | Section Title | Fatty Acids Identified, but C18:2 and C18:3 are not Present:
| 12.4.1.1 | Informational Text | Oil (or fat) with a low tendency toward self-heating was identified. This result does not eliminate the possibility that a substance with a high tendency toward self-heating was present prior to the fire.
Report Wording | 12.4.2 | Section Title | Fatty Acids Identified, C18:2 and C18:3 are Present but not the Most Abundant:
Report Wording | 12.4.2.1 | Informational Text | Oil (or fat) with a moderate tendency toward self-heating was identified. This result does not eliminate the possibility that a substance with a high tendency toward self-heating was present prior to the fire.
Report Wording | 12.4.3 | Section Title | Fatty Acids Identified, and C18:2 and C18:3 are Present and are the Most Abundant:
Report Wording | 12.4.3.1 | Informational Text | Oil (or fat) with a high tendency toward self-heating was identified. This does not mean that self-heating occurred nor can a probability be inferred that self-heating occurred.
Limitations | 13 | Section Title
Limitations | 13.1 | Recommendation | This is a qualitative technique and should not be used to quantitate FAME content.
Limitations | 13.2 | Informational Text | The identification of a specific type of oil requires a quantitative analysis of the triglycerides and is beyond the scope of this test method.
Limitations | 13.3 | Informational Text | There is no established error rate for the analysis of extracted and derivatized fatty acid-based products from fire debris and liquid samples by GC-MS.
Precision and Bias | 14 | Section Title
Precision and Bias | 14.1 | Informational Text | No information is presented about either the precision or bias of Test Method E2881 for the analysis and identification of FAMEs since the test result is nonquantitative.
ANNEX | A1 | Section Title | SAMPLE STORAGE
ANNEX | A1.1 | Requirement | After extraction, store the original debris or liquid sample using appropriate procedures for handling and documentation as set forth in Practice E1492.
ANNEX | A1.1.1 | Recommendation | Extract Storage Short-Term— Extracts prior to derivatization should be stored in a refrigerator in a sealed vial or container (free of extractable hydrocarbons) to inhibit further polymerization and prevent evaporation
ANNEX | A1.1.1.1 | Informational Text | Extracts after derivatization can be evaporated to dryness at room temperature and stored in a sealed vial or container (free of extractable hydrocarbons) at room temperature.
ANNEX | A1.1.2 | Recommendation | Extract Storage Long-Term— Only derivatized extracts should be stored long-term. The procedures set forth in Practice E2451 should be followed for long-term storage of extracts or liquid samples.
ANNEX | A1.1.3 | Requirement | Storage of ILR Extracts— Follow the procedures set forth in Practice E2451 for storage of ILR extracts.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed | Clause involves options; this option was not chosen | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Forensic Document Examination
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics