## Instructions for Use
AAFS Standards Resources and Training
This checklist is made possible through the following financial assistance award 70NANB24H078 awarded to AAFS from U.S. Department of Commerce, National Institute of Standards and Technology Each checklist is specific to a version of a standard, which is hyperlinked in the checklist. Not all sections of a standard are included in a checklist. A checklist is a supplementary tool and should not be used independently. To properly use a checklist, reference the complete standard. Go to https://www.nist.gov/organization-scientific-area-committees-forensic-science/access-standards for information on how to access published standards. A checklist does not provide interpretation guidance. If you have questions related to interpretation, please contact the standard developing organization (SDO). The SDO may have standard specific training available.
A checklist can be used for one or both of the following purposes: • to evaluate the implementation of a standard, or • to audit (internal or external) conformance with requirements in the standard and an FSSP’s management system
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions and an on-demand training webinar.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E2999-25a
ANSI/ASTM E2999-25a
Standard Test Methods for Analysis of Organic Compounds in Smokeless Powder by Gas Chromatography-Mass Spectrometry and Fourier Transform Infrared Spectroscopy
Discipline: | Fire Debris & Explosives
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Apparatus | 6 | Section Title
Apparatus | 6.1 | Requirement | Gas Chromatograph-Mass Spectrometer System— A gas chromatograph (GC) capable of using capillary columns and being interfaced to a mass spectrometer (MS) operating in electron ionization (EI) mode.
Apparatus | 6.1.1 | Requirement | Specimen Inlet System— An inlet system that can be operated in either split or splitless mode with capillary columns; the inlet system may use on-column technology.
Apparatus | 6.1.2 | Requirement | Column— A 30 m, 5 % Phenyl/95 % Methylpolysiloxane capillary column. Other column phases and lengths can be used provided that each component of the test mixture (see 7.3) is baseline separated.
Apparatus | 6.1.3 | Requirement | GC Oven— A column oven capable of reproducible temperature program operation in the range from 40 °C to 300 °C.
Apparatus | 6.1.4 | Requirement | Mass Spectrometer— A mass spectrometer capable of acquiring mass spectra from m/z 40 to m/z 400 with unit resolution or better, with continuous data output.
Apparatus | 6.1.5 | Requirement | Sensitivity— A system capable of detecting each compound in the test mixture (see 7.3) and providing sufficient ion intensity data to identify each compound, either by computer library search or by comparison with reference spectra.
Apparatus | 6.1.6 | Requirement | Fourier Transform Infrared Spectrometer— An FTIR capable of acquiring spectra in the mid-infrared region (wavenumber range of 4000 cm-1 to 650 cm-1).
Chemicals, Reagents, and Materials | 7 | Section Title
Chemicals, Reagents, and Materials | 7.1 | Requirement | Purity of Reagents— Reagent grade chemicals or better shall be used in all tests. Unless otherwise indicated, it is intended that all reagents conform to the specifications of the Committee on Analytical Reagents of the American Chemical Society where such specifications are available. Other grades may be used, provided it is first ascertained that the reagent is of sufficiently high purity to permit its use without lessening the accuracy of the determination.
Chemicals, Reagents, and Materials | 7.2 | Requirement | Solvents— Methylene chloride (DCM) and acetone. NOTE 1—Other solvents can be used if studies have demonstrated they can successfully extract the components in smokeless powder without adversely affecting the GC-MS or FTIR analysis.
Chemicals, Reagents, and Materials | 7.3 | Requirement | GC-MS Test Mixture— The test mixture consists of nitroglycerin, diphenylamine, ethyl centralite, and 2,6dinitrotoluene. The concentration of each compound is no greater than 0.005 % weight/volume (0.05 mg/mL) in the chosen solvent (see 7.2). Additional compounds commonly found in smokeless powders may also be included in the test mixture, such as methyl centralite, 2,4-dinitrotoluene, 2-nitrodiphenylamine, 4-nitrodiphenylamine, diethylphthalate, and dibutylphthalate. Refer to Table 3.
Chemicals, Reagents, and Materials | 7.3.1 | Informational Text | Similar concentrations of certified reference materials or individual standards of these compounds can be used in addition to or instead of a test mixture.
Chemicals, Reagents, and Materials | 7.4 | Informational Text | Internal Standard— An internal standard (for example, undecane, decane) may be used in the extraction solvent for GC-MS analysis.
Chemicals, Reagents, and Materials | 7.5 | Requirement | Smokeless Powders— Samples of smokeless powders, obtained as bulk reloading powders from commercial and retail sources or directly from the distributor or manufacturer, are used as reference samples.
Chemicals, Reagents, and Materials | 7.6 | Requirement | Polystyrene Film Standard.
Chemicals, Reagents, and Materials | 7.7 | Requirement | FTIR Supplies— Salt plates, mortar and pestle, pellet press.
Chemicals, Reagents, and Materials | 7.8 | Requirement | GC Carrier Gas— Helium or hydrogen of purity 99.995 % or higher.
Chemicals, Reagents, and Materials | 7.9 | Requirement | Glassware and Other Supplies— Disposable test tubes, pipettes, beakers, autosampler vials, weigh boats, weigh paper, watch glasses.
Quality Assurance and Quality Control | 8 | Section Title
Quality Assurance and Quality Control | 8.1 | Section Title | GC-MS:
Quality Assurance and Quality Control | 8.1.1 | Requirement | Prior to analyzing case specimens, evaluate the current tune results against acceptance criteria established in accordance with the instrument manufacturer’s recommendations, laboratory guidelines, or both. Tune the instrument if any acceptance criteria are deficient and retain the new tuning documentation as a portion of the quality control documentation.
Quality Assurance and Quality Control | 8.1.2 | Requirement | Analyze a negative control with every analytical batch to demonstrate that materials utilized in sample preparation and extraction, such as solvents, glassware and filtration apparatus are free from contaminants that could cause interferences.
Quality Assurance and Quality Control | 8.1.3 | Requirement | Analyze a prepared test mixture (see 7.3) or reference sample(s) at the beginning of each analytical batch to evaluate that instrument performance is acceptable.
Quality Assurance and Quality Control | 8.1.4 | Requirement | Maintain data from test mixtures or reference samples as a portion of casework documentation or quality control documentation, or both.
Quality Assurance and Quality Control | 8.1.5 | Requirement | Analyze a solvent blank between test mixtures, standards, reference and control samples, and questioned samples and evaluate solvent blank data to assess carryover or system contamination. NOTE 2—Analyzing a solvent blank between specimens is the preferred technique but is not required if studies demonstrate that the cleaning procedure is adequate to prevent carryover.
Quality Assurance and Quality Control | 8.2 | Section Title | FTIR:
Quality Assurance and Quality Control | 8.2.1 | Requirement | Prior to analyzing case specimens, check the FTIR’s most recent performance scan of a polystyrene film standard or equivalent reference sample against acceptance criteria in accordance with the manufacturer’s recommendation, laboratory guidelines, or both. Rescan if any acceptance criteria are deficient and retain documentation of the new performance check as a portion of the quality control documentation.
Quality Assurance and Quality Control | 8.2.1.1 | Requirement | Use commercially available standards or certified reference materials during performance checks.
Quality Assurance and Quality Control | 8.2.2 | Requirement | Perform a background scan prior to each specimen analysis and retain documentation in the case file.
Sample Handling | 9 | Section Title
Sample Handling | 9.1 | Requirement | Observe the appropriate procedures for handling and documentation of all submitted samples as described in Practices E1492, E3253, and E3255.
Sample Handling | 9.2 | Requirement | If an item is suspected of containing residues of an ignitable liquid, perform an ignitable liquid extraction and analysis on the item (or sample of bulk powder) prior to continuing with the analysis for smokeless powder. Refer to Guide E3245 and Test Method E1618.
Sample Handling | 9.2.1 | Requirement | Warning—Exceeding auto-ignition temperatures of suspected materials during headspace extraction could result in a fire.
Sample Handling | 9.3 | Informational Text | Refer to Practice E2998 for the analysis plan for the initial evaluation of the smokeless powder samples to determine the types of samples to be analyzed and practices and test methods to use.
Sample Handling | 9.3.1 | Informational Text | GC-MS analysis identifies the major organic compounds in smokeless powders with the exception of nitrocellulose and nitroguanidine.
Sample Handling | 9.3.2 | Informational Text | FTIR analysis is used in the identification of nitrocellulose and nitroguanidine on partial or whole grains of smokeless powder when sufficient sample is available.
Sample Handling | 9.3.2.1 | Informational Text | FTIR can also be used as an alternative technique to identify nitroglycerin. NOTE 3—FTIR analysis can include collection by attenuated total reflectance, transmission, or use of micro-FTIR.
Sample Handling | 9.4 | Informational Text | A smokeless powder sample(s) may be extracted and analyzed as a positive control(s) along with all other quality control samples referenced in 8.1.
Sample Preparation | 10 | Section Title
Sample Preparation | 10.1 | Section Title | Extraction A—Specimen Preparation for Single- and Double-base Powders:
Sample Preparation | 10.1.1 | Requirement | Caution: If limited sample is available use a single grain or partial grain for extraction.
Sample Preparation | 10.1.1.1 | Requirement | Place a sample of powder in a glass container, such as a test tube or vial, and extract with up to 300 μL of DCM. NOTE 4—The volume of the solvent can vary based on the amount, size, or shape of the powder and instrumental sensitivity.
Sample Preparation | 10.1.1.2 | Requirement | Extract for at least 30 min. Extraction times can vary depending upon grain morphology, size and surface coatings.
Sample Preparation | 10.1.1.3 | Requirement | Analyze the DCM extract for nitroglycerin and other common organic compounds in smokeless powders using Test Method 1.
Sample Preparation | 10.1.1.4 | Requirement | Remove and retain the remaining DCM extract in a cap-sealed vial for reanalysis if needed.
Sample Preparation | 10.1.1.5 | Requirement | Rinse the powder sample several more times with DCM.
Sample Preparation | 10.1.1.6 | Requirement | Remove any remaining DCM and allow the powder sample to dry. Take a portion of the dried sample and analyze for nitrocellulose by Test Method 2 (FTIR). NOTE 5—The clean-up procedure in Extraction C may be used prior to FTIR analysis to reduce spectral interferences from nitroglycerin.
Sample Preparation | 10.2 | Section Title | Extraction B—Specimen preparation for Triple-Base Powders:
Sample Preparation | 10.2.1 | Requirement | If nitroguanidine-like crystalline material is visually observed on a surface of the grain or partial grain, physically remove the material and analyze by Test Method 2 (FTIR). Proceed to Extraction A to extract and analyze other organic compounds in the sample.
Sample Preparation | 10.2.2 | Requirement | If crystalline material is not observed, slice open a single grain and examine visually, under a stereo microscope, or both.
Sample Preparation | 10.2.2.1 | Requirement | Use a microtome or other cutting device if the grain is difficult to slice open.
Sample Preparation | 10.2.2.2 | Requirement | If crystalline material is present, remove and analyze a portion by Test Method 2 (FTIR). Proceed to Extraction A to extract and analyze other organic compounds in the sample.
Sample Preparation | 10.2.2.3 | Requirement | If crystalline material is still not observed, analysis of nitroguanidine is not possible within the scope of this procedure. Proceed to Extraction A to extract and analyze other organic compounds in the sample.
Sample Preparation | 10.3 | Section Title | Extraction C—Optional Clean-up for FTIR of Nitrocellulose:
Sample Preparation | 10.3.1 | Requirement | Dissolve dry material in the minimal amount of acetone needed to cover the material.
Sample Preparation | 10.3.2 | Requirement | Disperse the dissolved material as a thin layer on a glass slide or watch glass, or in another suitable glass container.
Sample Preparation | 10.3.3 | Requirement | Allow the dispersed material to dry, at room temperature, to a white solid material.
Sample Preparation | 10.3.4 | Requirement | Rinse the dried residue repeatedly with approximately 500 μL to 1 mL aliquots of DCM to remove the other organic compounds. The volume and number of aliquots will vary with the amount of solid residue.
Sample Preparation | 10.3.5 | Requirement | Allow the material to dry then analyze the residue by Test Method 2 (FTIR).
GC-MS Analysis (Test Method 1) | 11 | Section Title
GC-MS Analysis (Test Method 1) | 11.1 | Requirement | Analyze DCM extracts for common smokeless powder compounds using the GC-MS conditions listed in Table 1.
GC-MS Analysis (Test Method 1) | 11.1.1 | Informational Text | Conditions other than those listed in Table 1 can be used if they have been validated or verified and provide sufficient sensitivity and resolution to separate the compounds in the test mixture.
GC-MS Conditions | Table 1 | Section Title
GC-MS Conditions | Table 1 | Requirement | GC column: 30 meter, DB-5 capillary column
GC-MS Conditions | Table 1 | Requirement | Initial temp: 45°C
GC-MS Conditions | Table 1 | Requirement | Initial hold time: 3 minutes
GC-MS Conditions | Table 1 | Requirement | Ramp 1: 15 °C ⁄minute to 150 °C
GC-MS Conditions | Table 1 | Requirement | Ramp 2: 40 °C ⁄minute to 265 °C
GC-MS Conditions | Table 1 | Requirement | Final hold time: 7 minutes
GC-MS Conditions | Table 1 | Requirement | Injector temperature: 190 °C to 220 °C
GC-MS Conditions | Table 1 | Requirement | Mode: Electron Ionization (EI)
GC-MS Conditions | Table 1 | Requirement | MS scan range: 41 to 400 m/z
GC-MS Conditions | Table 1 | Requirement | Solvent delay: 3 minutes
GC-MS Analysis (Test Method 1) | 11.2 | Informational Text | A batch analysis sequence consists of a smokeless powder test mixture or individual standards, solvent blanks, negative controls, and questioned samples. Additional samples, such as positive controls and reference samples, may also be included.
GC-MS Analysis (Test Method 1) | 11.2.1 | Informational Text | Analyzing a solvent blank between samples is recommended but is not necessary if studies demonstrate that the cleaning procedure is adequate to prevent carryover.
FTIR Analysis (Test Method 2) | 12 | Section Title
FTIR Analysis (Test Method 2) | 12.1 | Requirement | Analyze specimens using the FTIR parameters listed in Table 2.
FTIR Parameters | Table 2 | Section Title
FTIR Parameters | Table 2 | Requirement | Minimum Wavenumber range: 4000 to 650 cm1
FTIR Parameters | Table 2 | Requirement | Minimum Background scans: 32
FTIR Parameters | Table 2 | Requirement | Minimum Sample scans: 32
FTIR Parameters | Table 2 | Requirement | Resolution: 4 cm-1
FTIR Analysis (Test Method 2) | 12.2 | Requirement | Clean the ATR stage or window with an appropriate solvent before all background and specimen scans.
FTIR Analysis (Test Method 2) | 12.3 | Requirement | Perform a background scan prior to each specimen analysis.
FTIR Analysis (Test Method 2) | 12.4 | Requirement | If FTIR results are skewed or otherwise unacceptable for a questioned sample, repeat the analysis with a new sample if additional material is available (see 14.3).
FTIR Analysis (Test Method 2) | 12.4.1 | Requirement | Nitroglycerin and nitrocellulose have similar FTIR spectral features. Use Extractions A and C (see 10.1 and 10.3) to minimize spectral interferences.
Data Processing and Interpretation | 13 | Section Title
Data Processing and Interpretation | 13.1 | Section Title | General Data Processing Scheme:
Data Processing and Interpretation | 13.1.1 | Informational Text | GC-MS uses the total ion chromatogram (TIC) and extracted ion profiling (EIP) to identify the compounds of interest based on retention times and mass spectral data.
Data Processing and Interpretation | 13.1.2 | Informational Text | FTIR analysis is used for the identification of nitrocellulose and nitroguanidine.
Data Processing and Interpretation | 13.1.2.1 | Informational Text | Nitroglycerin and other organic compounds can also be identified by FTIR.
Data Processing and Interpretation | 13.2 | Section Title | Processing and Evaluation of GC-MS Data:
Data Processing and Interpretation | 13.2.1 | Requirement | Use extracted ion profiles (EIPs) to generate mass spectra of compounds of interest for each questioned sample, positive control, test mixture, and reference sample analyzed. Compounds of interest and their associated ions are listed in Table 3. NOTE 6—Additional components that are either minor compounds or are not commonly seen in the analysis of smokeless powders do not appear in the table. Manufacturers can change suppliers or introduce new products or formulations with additional compounds that do not appear in Table 3.
Key m/z Ions for Smokeless Powders | Table 3 | Section Title | Compound; m/z
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | Ethyleneglycol dinitrateA; 46, 76
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | Nitroglycerin; 46, 76
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | CamphorA; 95, 152
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | 2,4- and 2,6-dinitrotoluene; 165, 182
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | 2,4,6-trinitrotoluene (TNT)A; 210
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | Ethylcentralite; 120, 268
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | MethylcentraliteA; 134, 240
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | Diphenylamine; 169
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | 2- or 4-nitrodiphenylamine; 167, 214
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | 2- or 4-nitrosodiphenylamine; 167, 198
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | 1-methyl-3,3-diphenylurea (AkarditeII)A; 169, 226
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | Dibutyl phthalate; 149, 223
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | Diisoamyl phthalateA; 149
Key m/z Ions for Smokeless Powders | Table 3 | Requirement | Dioctyl phthalateA; 149, 279
Key m/z Ions for Smokeless Powders | Table 3 | Informational Text | A Denotes a major component but present in only a limited number of products.
Data Processing and Interpretation | 13.2.2 | Requirement | Evaluate peaks in the TIC and EIPs of each questioned sample to identify the compounds of interest based on retention times and mass spectral data in comparison to the data from test mixture(s), certified reference materials, or individual standards (see 7.3 and 7.3.1).
Data Processing and Interpretation | 13.2.3 | Requirement | For solvent blanks and negative controls, confirm that there are no peaks present in the TIC that correspond to compounds of interest.
Data Processing and Interpretation | 13.2.4 | Requirement | Retain chromatograms and mass spectra for test mixtures (or individual standards), reference samples, questioned samples, solvent blanks, and control samples in the case file.
Data Processing and Interpretation | 13.3 | Section Title | FTIR Data Evaluation:
Data Processing and Interpretation | 13.3.1 | Requirement | Use a reference spectral library to search the FTIR spectrum of a questioned specimen for possible matches or compare the questioned spectrum to the spectrum of a reference sample or standards. Compare the spectral bands of the library matches or reference sample to the questioned sample spectral bands for identification of a compound.
Data Processing and Interpretation | 13.3.2 | Requirement | Retain the FTR spectra and results of the comparison in the case file.
Data Processing and Interpretation | 13.3.3 | Informational Text | A mixture of compounds may be present in DCM extracts of smokeless powder.
Data Processing and Interpretation | 13.3.3.1 | Requirement | Use multi-component search software to separate peaks in mixtures.
Data Processing and Interpretation | 13.3.3.2 | Requirement | Visually confirm the agreement of the composite spectrum generated from the multi-component search with the original spectrum.
Data Processing and Interpretation | 13.3.3.3 | Requirement | Retain the FTIR spectrum, the multi-component search results, and the comparison of these results to the original spectrum in the case file.
Data Processing and Interpretation | 13.4 | Section Title | Interpretation of Results:
Data Processing and Interpretation | 13.4.1 | Informational Text | GC-MS and FTIR analysis data is used in conjunction with the physical characteristics of powder grains as the basis for identification or classification of a smokeless powder. Refer to Practice E2998 for criteria for the identification and classification of smokeless powders.
General Considerations and Limitations | 14 | Section Title
General Considerations and Limitations | 14.1 | Informational Text | Some organic compounds, in addition to nitrocellulose and nitroguanidine, that are associated with smokeless powders, are not detectable by these test methods (for example, rosin, graphite).
General Considerations and Limitations | 14.2 | Informational Text | Some compounds in smokeless powders are proprietary materials. The unavailability of a standard or equivalent reference sample for comparison can hinder the analyst’s ability to identify such compounds.
General Considerations and Limitations | 14.3 | Recommendation | Some smokeless powders contain a small percentage of grains with a high concentration of an inorganic salt (flash reducer or burn modifier). Skewed FTIR results could occur if one of these grains was randomly chosen and analyzed by FTIR. The analyst should repeat the FTIR analysis with a new sample in such instances where possible.
General Considerations and Limitations | 14.4 | Informational Text | When a mixture of compounds is present in a DCM extract, the identification of minor compounds might not be possible with FTIR.
Report Writing | 15 | Section Title
Report Writing | 15.1 | Informational Text | Refer to Practice E620 for general information on report writing.
Report Writing | 15.2 | Informational Text | Refer to Practice E2998 for reporting results for the analysis of smokeless powders.
Precision and Bias | 16 | Section Title
Precision and Bias | 16.1 | Informational Text | These test methods are qualitative with no numerical results, and therefore measures of precision and bias are not applicable.
Precision and Bias | 16.2 | Recommendation | Other factors to consider that are relevant to the acceptability of test results include sensitivity and specificity, selectivity, interferences, repeatability and reproducibility, instrumental detection limits, and user subjectivity. These factors should be considered and evaluated during in situ method validation.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed | Clause involves options; this option was not chosen | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Forensic Document Examination
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics