## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ASTM E3255-21
ASTM E3255-21
Standard Practice for Quality Assurance of Forensic Science Service Providers Performing Forensic Chemical Analysis
Discipline: | Interdisciplinary
Checklist version: | V1.1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Quality Management System | 5 | Section Title
Quality Management System | 5.1 | Informational Text | It is the goal of a forensic science service provider’s evidence analysis program to provide customers with high quality analyses that produce reliable and accurate results.
Quality Management System | 5.2 | Requirement | A documented quality management system shall be established and maintained.
Quality Management System | 5.2.1 | Requirement | Personnel responsible for the quality management system (that is, quality assurance manager) shall be clearly designated.
Quality Management System | 5.2.2 | Requirement | The quality management system shall cover all procedures and reports associated with analysis.
Quality Management System | 5.2.3 | Requirement | The quality management system shall conform to an international standard, such as ISO/IEC 17025 or ISO/IEC 17020.
Personnel | 6 | Section Title
Personnel | 6.1 | Requirement | Personnel shall have defined responsibilities, duties, educational requirements and required skills.
Personnel | 6.2 | Requirement | Forensic science service providers shall conform to Practice E2917.
Personnel | 6.3 | Requirement | Designated key personnel and responsibilities shall be established and maintained. An individual (however named) can be responsible for more than one of the duties detailed in 6.4 to 6.9.
Director | 6.4 | Informational Text | A designated person who is responsible for the overall operation and administration of the forensic science service provider, including the employment and training of personnel and assuring compliance with applicable regulations.
Quality Assurance Manager | 6.5 | Informational Text | A designated person who is responsible for maintaining the quality management system (including an annual review of the program) and who ensures conformance with the program.
Technical Support Personnel | 6.6 | Requirement | A person who performs basic forensic science duties (for example, calibration checks, making solutions, glassware washing, etc.), but does not analyze evidence. Technical support personnel shall:
Technical Support Personnel | 6.6.1 | Requirement | Have education, skills, and abilities commensurate with their responsibilities; and
Technical Support Personnel | 6.6.2 | Requirement | Have on-the-job training specific to the responsibilities of their position.
Technician | 6.7 | Requirement | A person who analyzes evidence, but does not issue reports of analytical results or interpretations. Technicians shall:
Technician | 6.7.1 | Requirement | Have education, skills, and abilities commensurate with their responsibilities; and
Technician | 6.7.2 | Requirement | Have formal training in evidence handling and those aspects of their duties specific to analysis. These include: (1) Successful and documented completion of written or oral examinations, (2) Successful and documented completion of competency testing, and (3) Successful and documented completion of a moot/mock court exercise.
Forensic Science Practitioner | 6.8 | Requirement | A designated person who may be authorized to: examine and analyze materials, or direct such examinations to be performed; interpret data, issue reports for court or investigative purposes; and conduct, method validation, training, and technical review of reports. Forensic science practitioners shall:
Forensic Science Practitioner | 6.8.1 | Requirement | Meet all the requirements of a technician (6.7), and
Forensic Science Practitioner | 6.8.2 | Requirement | Have at least a bachelor’s degree or equivalent (generally, a four-year postsecondary degree) in a natural or physical science. The individual shall have successfully completed lecture and laboratory classes’ in chemistry specific to their particular discipline, and
Forensic Science Practitioner | 6.8.3 | Requirement | Conform to Practice E2917.
Technical Leader | 6.9 | Requirement | A designated person or team of individuals who have the overall responsibility and authority for the technical operations associated with evidence analyses. Technical operations include maintaining protocols, developing analytical methodology, conducting method validation, and performing technical review of reports. The technical leader shall:
Technical Leader | 6.9.1 | Requirement | Meet all the requirements of a forensic science practitioner (6.8), and
Technical Leader | 6.9.2 | Requirement | Have a minimum of three (3) years of experience as a qualified forensic science practitioner performing independent evidence analyses in the specific discipline.
Physical Plant | 7 | Section Title
Physical Plant | 7.1 | Requirement | Forensic science service providers shall provide:
Physical Plant | 7.1.1 | Requirement | Adequate safety and security for personnel and operations;
Physical Plant | 7.1.2 | Requirement | Facilities that meet required health and safety standards and applicable building codes;
Physical Plant | 7.1.3 | Requirement | Suitable space to perform required analytical functions and prevent contamination;
Physical Plant | 7.1.4 | Requirement | Environmental and procedural controls to prevent incidental contamination;
Physical Plant | 7.1.5 | Requirement | Engineering devices and personal protective equipment (for example, chemical fume hoods, safety showers, gloves, safety glasses, etc.) to protect personnel from chemical hazards; and
Physical Plant | 7.1.6 | Requirement | Facilities that ensure safe and secure storage of evidence, chemicals, reagents, reference materials and records.
Evidence Control | 8 | Section Title
Evidence Control | 8.1 | Requirement | Forensic science service providers shall follow a documented control system to ensure the integrity of evidence.
Evidence Control | 8.2 | Informational Text | See Guide E1459 for physical evidence labeling and related documentation.
Evidence Control | 8.3 | Informational Text | See Practice E1492 for receiving, documenting, storing, and retrieving evidence.
Integrity of Evidence | 8.4 | Requirement | Evidence shall be properly secured and sealed. Appropriate storage conditions shall ensure that, insofar as possible, the composition of evidentiary material is not altered. The integrity of the evidence shall be protected in ways that safeguard against loss, deterioration or contamination. Alteration of the evidence or packaging shall be documented.
Accounting for Evidence | 8.5 | Requirement | Evidence that is maintained by the forensic science service provider shall be inventoried, at a minimum, annually.
Disposition of Evidence | 8.6 | Requirement | Records shall be kept regarding the disposition of all items of evidence (see Practice E1492).
Method Validation and Verification | 9 | Section Title
Method Validation and Verification | 9.1 | Requirement | Validation is required to demonstrate that methods are suitable for their intended purpose.
Method Validation and Verification | 9.2 | Requirement | Forensic science service providers adopting methods validated elsewhere shall verify the validity and reliability of those methods used in their environment, with their instrumentation, equipment, reagents and chemicals, and forensic science practitioners prior to use.
Method Validation and Verification | 9.3 | Requirement | New methods developed for characterization, identification, and comparison of chemicals shall be based on accepted scientific principles. The forensic science service provider shall perform validation studies to establish both the method’s validity and reliability prior to use in casework.
Method Validation and Verification | 9.3.1 | Requirement | Minimum acceptance criteria shall be established for method performance and be described along with means for demonstrating compliance.
Method Validation and Verification | 9.4 | Requirement | Method validation and verifications shall address selectivity, precision, and other relevant performance characteristics and limitations.
Method Validation and Verification | 9.4.1 | Requirement | Validations and verifications involving quantitative analysis shall also address limits of detections, lower limit of quantitation, linearity, accuracy, and uncertainty of the measurement process.
Method Validation and Verification | 9.5 | Requirement | Validation and verifications shall be approved by the technical leader or other designated personnel prior to use in casework.
Method Validation and Verification | 9.6 | Requirement | Validation and verification data and documentation shall be maintained.
Analytical Procedures | 10 | Section Title
Analytical Procedures: | 10.1 | Section Title
Analytical Procedures: | 10.1.1 | Requirement | Forensic science service providers shall follow documented, validated, and authorized analytical methods.
Analytical Procedures: | 10.1.1.1 | Requirement | The technical leader or other designated individual shall be responsible for evaluation and authorization of the analytical procedures used by the forensic science service provider, including reviews of method validations and verifications. All evaluations and authorizations shall be documented and retained.
Analytical Procedures: | 10.1.1.2 | Requirement | A chemical/material identification, quantification or comparison shall be based on specific characteristics defined in the validated method, and decisions on the type/source of the material should be based on pre-established decision-thresholds.
Analytical Procedures: | 10.1.1.3 | Requirement | Relevant limitations of chemical/material identifications or comparisons shall be evaluated and reported.
Analytical Procedures: | 10.1.1.4 | Requirement | Forensic science service providers shall have in place protocols for the sampling of evidence.
Analytical Procedures: | 10.1.1.5 | Recommendation | When applicable, forensic science service providers should employ a statistically based protocol when an inference is to be made regarding a specified proportion of a submission for qualitative or quantitative determinations.
Analytical Procedures: | 10.1.2 | Recommendation | Published or standardized methods are recommended. Laboratory-developed or modified methods can also be used.
Analytical Procedures: | 10.1.2.1 | Requirement | The forensic science service provider shall verify published or standardized methods by in-house performance testing.
Analytical Procedures: | 10.1.2.2 | Requirement | The use of laboratory-developed or modified methods shall be validated and approved by the technical leader or other designated individual according to forensic science service provider policy.
Analytical Procedures: | 10.1.3 | Requirement | Work practices shall be established to prevent contamination of evidence during analysis.
Analytical Procedures: | 10.1.3.1 | Requirement | The processing of trace/residue samples should be separated in space from bulk evidence submissions to prevent incidental contamination. If space does not allow for this, then the processing of bulk and trace/residue samples shall be separated by time with thorough cleaning between processing steps.
Analytical Procedures: | 10.1.3.2 | Requirement | Forensic science practitioners shall take measures to be assured that identifications are correct (for example, excluding contamination, instrument carryover, etc.) and relate to the correct submission.
Analytical Procedures: | 10.1.4 | Requirement | Forensic science service providers shall follow documented procedures for the evaluation and acceptance of data based on the method utilized.
Analytical Procedures: | 10.1.5 | Requirement | Forensic science service providers shall monitor analytical processes using appropriate blanks, positive and negative controls and reference materials.
Analytical Procedures: | 10.1.5.1 | Requirement | Negative controls shall be run concurrently with evidence to demonstrate that the sampling devices, chemicals, instruments and extraction processes do not result in contamination of the evidence.
Analytical Procedures: | 10.1.5.2 | Requirement | The forensic science service provider shall have a procedure for routinely testing the reliability of the analytical processes with known reference materials (positive control). This can be done concurrently with each analysis or on a predefined schedule.
Analytical Procedures: | 10.1.6 | Requirement | Reference materials and associated data shall be used to demonstrate the reliability of test results. A positive test result shall meet the acceptance criteria defined in the method validation and operating protocol. In descending order of preference, the acceptance criteria should be based on:
Analytical Procedures: | 10.1.6.1 | Requirement | Comparison of instrumental test/case sample data to suitable reference material data analyzed under the same analytical conditions (that is, instrument, instrumental methods, and stationary phase). If reference material data is collected on a different instrument than the test/case sample, it shall be demonstrated that both instruments produce comparable data. The reference material may be analyzed: (1) Contemporaneously with test/case sample (for example, same sequence/batch); (2) As part of routine quality control (for example, daily check solutions); or (3) At a previous date (for example, method validation, internal reference collection).
Analytical Procedures: | 10.1.6.2 | Requirement | Comparison of instrumental test/case sample data to external reference data when a reference material is unavailable. External reference data shall be assessed and demonstrated to be fit for purpose. Factors include: (1) Origin of the data; (2) Validation of the data; (3) Peer review of the data; and (4) Comparability of analytical conditions.
Analytical Procedures: | 10.1.6.3 | Recommendation | Structural elucidation techniques, when applicable by discipline, and when neither reference materials nor external reference data are available. Structural elucidation techniques may be employed by forensic science practitioners that are competent in structural elucidation interpretation.
Assessment of Reference Materials: | 10.2 | Section Title
Assessment of Reference Materials: | 10.2.1 | Requirement | Forensic science service providers shall have a process for assessing that reference materials are fit for purpose.
Assessment of Reference Materials: | 10.2.1.1 | Informational Text | Assessments are not required for reference materials obtained from a provider accredited under ISO 17034.
Assessment of Reference Materials: | 10.2.1.2 | Requirement | The assessment and purpose (for example, qualitative or quantitative) of a reference material shall be documented. The documentation shall include: identity; source; assigned unique identifier; date and name of the individual who performed the assessment; and verification test data.
Assessment of Reference Materials: | 10.2.1.3 | Requirement | To be fit for purpose, reference materials shall meet minimum specifications defined in the forensic science service provider’s validated methods.
Assessment of Reference Materials: | 10.2.1.4 | Requirement | The assessment shall be performed on each lot of reference materials.
Assessment of Reference Materials: | 10.2.1.5 | Requirement | The assessment shall be completed prior to casework analysis.
Assessment of Reference Materials: | 10.2.2 | Requirement | Reference materials shall only be used for the purpose defined by the forensic science service provider. For example, a reference material may be deemed suitable for qualitative but not quantitative determinations.
Assessment of Reference Materials: | 10.2.2.1 | Requirement | Fit for purpose for qualitative work requires an assessment of chemical identity. Examples of verification of chemical identity by analysis include: (1) Comparison of the results to peer-reviewed published data, (2) Data produced by an accredited forensic science service provider, (3) Data produced from a previously verified reference material, or (4) Evaluation of data from in-house structural elucidation analysis of the material.
Assessment of Reference Materials: | 10.2.2.2 | Requirement | Fit for purpose for quantitative work requires traceability and assessment of purity or concentration, or both, as appropriate to the application and its associated uncertainty of measurement in addition to 10.2.3. Examples of verification of purity by analysis utilizing validated methods include: (1) Quantitative nuclear magnetic resonance spectroscopy, (2) Quantitative ultraviolet-visible spectroscopy, (3) Quantitative emission spectroscopy, or (4) Comparison to previously verified material.
Assessment of Reference Materials: | 10.2.2.3 | Recommendation | For quantitative determinations, different sources of reference material should be used for calibration and quality control. Where this is not feasible, two different lots of the same source may be used or lastly a single source of reference material can be subdivided and each part assigned a specific purpose.
Assessment of Reference Materials: | 10.2.2.4 | Recommendation | Certified reference materials are preferred, but not required, as they are not typically available for many chemicals.
Assessment of Reference Materials: | 10.2.3 | Informational Text | The specifications in 10.2 can be described in a certificate, statement of analysis, data sheet or label supplied with the material or can be determined by in-house analysis or reference to published literature.
Assessment of Reference Materials: | 10.2.4 | Requirement | The forensic science service provider shall assess the veracity of the information supplied with a reference material even if the material meets the definition of a certified reference material.
Assessment of Reference Materials: | 10.2.5 | Requirement | Reference materials that degrade or otherwise change in composition or concentration over time shall have an expiration date.
Assessment of Reference Materials: | 10.2.5.1 | Recommendation | If the material is not supplied with an expiration date, one should be assigned at the first assessment (10.2.1).
Assessment of Reference Materials: | 10.2.5.2 | Requirement | If the expiration date passes before the material is fully used, then the material can be re-assessed and the expiration date extended. The forensic science service provider protocol for extending expiration dates shall be documented and include analysis of the material.
Instrument/Equipment Performance | 11 | Section Title
Instruments | 11.1 | Requirement | A performance check of each instrument shall be conducted prior to being initially used for casework. This performance check and the data collected shall be documented and include:
Instruments | 11.1.1 | Requirement | Analysis and comparison of known materials to include expected analytes, in order to demonstrate its suitability for intended use.
Instruments | 11.1.2 | Requirement | Comparison of data created to that of comparable published data.
Instrument Performance | 11.2 | Requirement | Instruments shall be routinely optimized and monitored to ensure that proper performance is maintained.
Instrument Performance | 11.2.1 | Requirement | Acceptance criteria for monitoring instrument performance shall be defined and documented.
Instrument Performance | 11.2.2 | Requirement | Monitoring shall include, at a minimum, the use of reference materials, test mixtures and blanks when applicable.
Instrument Performance | 11.2.3 | Requirement | Instrument performance monitoring and the data collected shall be documented and retained.
Instrument Performance | 11.2.4 | Requirement | The manufacturer’s operation manual and other relevant documentation for each instrument should be readily available.
Equipment | 11.3 | Requirement | Only suitable and properly working equipment shall be used in the course of analyses.
Equipment | 11.3.1 | Recommendation | Equipment performance parameters should be routinely monitored and documented when applicable.
Equipment | 11.3.2 | Recommendation | The manufacturer’s operation manual and other relevant documentation for each piece of equipment should be readily accessible.
Chemicals and Reagents | 12 | Section Title
Chemicals and Reagents | 12.1 | Requirement | Chemicals and reagents used in forensic chemical analysis shall be at least ACS reagent grade.
Chemicals and Reagents | 12.2 | Requirement | There shall be documented procedures for the formulation of all chemical reagents produced.
Chemicals and Reagents | 12.2.1 | Requirement | Documentation for reagents prepared shall include identity, concentration (when appropriate), date of preparation, identity of the individual preparing the reagents, storage conditions (if appropriate), and the expiration date (if appropriate).
Chemicals and Reagents | 12.3 | Requirement | The efficacy of all reagents shall be checked prior to or concurrent with their use in casework. Results of these tests shall be documented and retained.
Chemicals and Reagents | 12.4 | Requirement | The received and opened date(s) shall be recorded for chemicals and reagents, where relevant to testing results.
Chemicals and Reagents | 12.5 | Requirement | Chemical and reagent containers shall be labeled as to their contents.
Casework Documentation, Report Writing, and Review | 13 | Section Title
Casework Documentation: | 13.1 | Section Title
Casework Documentation: | 13.1.1 | Requirement | Maintain documentation that contains sufficient information and clarity to allow another properly trained forensic science practitioner the ability to evaluate the notes, interpret data, and verify whether the original result is accurate.
Casework Documentation: | 13.1.2 | Requirement | Evidence handling documentation shall establish chain-of-custody.
Casework Documentation: | 13.1.3 | Requirement | Case related communications shall be maintained.
Casework Documentation: | 13.1.4 | Requirement | Analytical documentation shall include observations, methodology, blanks, controls, test results, and supporting documentation. Examples of supporting documentation are instrumental data, sequence files or run logs, charts, graphs, and spectra generated during analysis.
Casework Documentation: | 13.1.4.1 | Requirement | Any deviations from required analytical procedures shall be documented (see 10.1.1).
Casework Documentation: | 13.1.5 | Requirement | Casework documentation shall be retained according to the forensic science service provider’s policy or jurisdictional regulations or laws.
Case Review: | 13.2 | Section Title
Case Review: | 13.2.1 | Requirement | Forensic science service providers shall have a documented policy for reviewing and authorizing reported results.
Case Review: | 13.2.2 | Recommendation | When achievable, reviewers should not be privy to analysts’ results when conducting reviews.
Case Review: | 13.2.3 | Requirement | Forensic science service providers shall have a documented procedure for resolving instances where forensic science practitioner and reviewer disagree. The procedure shall include instruction to document the disagreement and corresponding resolution.
Case Review: | 13.2.4 | Requirement | Identified deficiencies shall be resolved in accordance with Section 16.
Report Writing: | 13.3 | Section Title
Report Writing: | 13.3.1 | Requirement | Reports shall include relevant elements described in Practice E620 and ISO/IEC 17025.
Report Writing: | 13.3.2 | Requirement | Reports issued by the forensic science service provider shall be accurate, clear, objective, and meet the requirements of the jurisdictions served.
Competency and Proficiency Testing | 14 | Section Title
Competency and Proficiency Testing | 14.1 | Requirement | Forensic science service providers shall establish a documented competency testing and proficiency testing program. Forensic science service providers shall have documented protocols for evaluating and monitoring the competency (see Practice E2917) and continuing proficiency of its forensic science practitioners.
Competency Testing: | 14.2 | Section Title
Competency Testing: | 14.2.1 | Requirement | Forensic science service providers shall test the competency of their technicians and forensic science practitioners prior to assigning them independent casework responsibilities.
Competency Testing: | 14.2.2 | Recommendation | Competency test samples should be representative of the forensic science service provider’s range of casework.
Competency Testing: | 14.2.3 | Requirement | Methodology utilized to perform competency tests shall be the same as those practiced when analyzing evidence.
Proficiency Testing: | 14.3 | Section Title
Proficiency Testing: | 14.3.1 | Requirement | Forensic science service providers shall participate in proficiency testing in order to evaluate the continuing capability and performance of their procedures and persons engaged in testing.
Proficiency Testing: | 14.3.1.1 | Requirement | Performance shall be evaluated against preestablished criteria.
Proficiency Testing: | 14.3.1.2 | Requirement | Each person engaged in testing proficiency samples shall not be privy to the known composition. When achievable, forensic science service providers should implement blind proficiency testing programs.
Proficiency Testing: | 14.3.1.3 | Requirement | Each person engaged in testing activities related to evidence analysis shall successfully participate in at least one intralaboratory comparison, interlaboratory comparison, or proficiency test per calendar year that addresses the specific forensic chemistry discipline or category of testing for which they are authorized.
Proficiency Testing: | 14.3.1.4 | Requirement | When available, each forensic science service provider engaged in testing activities related to evidence analysis shall participate in at least one proficiency test each calendar year that addresses the specific forensic chemical discipline or category of testing. The proficiency test provider should conform to ISO/IEC 17043.
Proficiency Testing: | 14.3.2 | Recommendation | Proficiency test samples should be representative of the forensic science service provider’s range of casework.
Proficiency Testing: | 14.3.3 | Requirement | Methodology utilized to perform proficiency tests shall be the same as those normally used when analyzing evidence.
Forensic Science Service Provider Audits | 15 | Section Title
Forensic Science Service Provider Audits | 15.1 | Requirement | Internal audits of the forensic science service provider’s operations shall be conducted at least once a year. Record keeping shall satisfy the requirements of ISO/IEC 17025.
Nonconforming Work | 16 | Section Title
Nonconforming Work | 16.1 | Requirement | In the course of examining evidence, forensic science service providers can expect to encounter some operations or results that do not conform in some manner. Examples of nonconforming work include: an erroneous analytical result or interpretation; or any unapproved deviation from an established policy or procedure in an analysis. Nonconforming work shall satisfy the requirements of ISO/IEC 17025. NOTE 1—It should be recognized that to be effective, the definition for “nonconforming work” must be relatively broad. As such, nonconforming work can have markedly different degrees of seriousness. For example, a misidentification of a compound or mixture would be very serious and perhaps require that either the methodology or the forensic science practitioner be removed from casework pending appropriate remedial action, as determined by the forensic science service provider. However, other nonconforming work could be more clerical in nature, requiring a simple correction without any suspension of methodology or reassignment of personnel. Thus, it can be advantageous to identify the differing levels of seriousness for nonconforming work and make the action required commensurate with the seriousness.
Nonconforming Work | 16.1.1 | Requirement | When nonconforming work is encountered or discovered, the cause and the impact on past work shall be determined. Corrective measures shall be documented to minimize the likelihood of recurrence.
Summary of Required Documentation | 17 | Section Title
Summary of Required Documentation | 17.1 | Requirement | In addition to casework documentation, forensic science service providers shall maintain records on the following topics:
Summary of Required Documentation | 17.1.1 | Requirement | Test methods/procedures for evidence analysis;
Summary of Required Documentation | 17.1.1.1 | Requirement | Validation and verification data and results, and
Summary of Required Documentation | 17.1.1.2 | Requirement | Statistical performance of test methods (including uncertainty, population inferences, control charts, etc.).
Summary of Required Documentation | 17.1.2 | Requirement | Reference materials;
Summary of Required Documentation | 17.1.3 | Requirement | Preparation and checks of reagents;
Summary of Required Documentation | 17.1.4 | Requirement | Evidence handling protocols;
Summary of Required Documentation | 17.1.5 | Requirement | Instrument calibration and maintenance;
Summary of Required Documentation | 17.1.6 | Requirement | Instrument inventory (for example, manufacturer, model, serial number, acquisition date);
Summary of Required Documentation | 17.1.7 | Requirement | Competency and proficiency testing;
Summary of Required Documentation | 17.1.8 | Requirement | Personnel training and qualification;
Summary of Required Documentation | 17.1.9 | Requirement | Quality assurance protocols and audits; and
Summary of Required Documentation | 17.1.10 | Requirement | Health, safety, and security protocols.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Bloodstain Pattern Analysis | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Crime Scene Investigation & Reconstruction | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Digital Evidence | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Dogs & Sensors | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Facial Identification | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Fire & Explosion Investigation | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Firearms & Toolmarks | Informational Text cont'd
Footwear & Tire
Forensic Anthropology
Forensic Document Examination
Forensic Nursing
Forensic Odontology
Forensic Toxicology
Friction Ridge
Human Forensic Biology
Ignitable Liquids, Explosives, & Gunshot Residue
Interdisciplinary
Medicolegal Death Investigation
Seized Drugs
Speaker Recognition
Trace Materials
Video/Imaging Technology & Analysis
Wildlife Forensic Biology