## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASB 036 2019-1st Ed
ANSI/ASB 036 - 2019 1st Edition
Standard Practices for Method Validation in Forensic Toxicology
Discipline: | Toxicology
Checklist version: | V1.2
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
When to Validate Methods | 4 | Section Title
When to Validate Methods | 4 | Requirement | Methods shall be validated to verify a method's performance parameters are fit for use for a particular analysis. Common examples include: a) new analytical method; b) modifications of an established analytical method to improve performance or extend its use beyond that for which it was originally validated (e.g., addition of new compounds to the method’s scope); c) to demonstrate equivalence between an established method/instrument and a new method/instrument; d) existing analytical methods that do not currently meet the requirements of this document. The parameters to be evaluated for validation of methods will depend upon the circumstances in which the method is to be used. Likewise, it is recognized that after validation has occurred, methods may be revised. The extent and frequency of revalidation of previously validated methods will depend upon the nature of the intended changes or laboratory policy. See Section 10 for further guidance on revalidation of previously validated methods. Laboratories using methods that were validated prior to the promulgation of this document shall demonstrate and document that those methods are fit-for-purpose under this standard. These methods will likely have sufficient historical calibration and control data, as well as previously analyzed casework sample results, that can be used to address a number of the required validation parameters. In the absence of sufficient data to fulfill these minimum standards, appropriate studies shall be conducted to ensure compliance with this document.
Method Development and Optimization | 5 | Section Title
General | 5.1 | Requirement | For purposes of this document, method development will be considered in two phases: 1) instrumental and data acquisition/processing parameters and 2) sample preparation. It is essential that validation is conducted with the same analytical conditions and techniques as the final developed method. The principles of good laboratory practice and record keeping shall be applied to the concepts of this document. This includes documentation of parameters that were evaluated during method development, yet did not provide acceptable results.
Development and Optimization of Instrumental and Data Processing Parameters | 5.2 | Requirement | Instrumental and data processing parameters are defined and optimized through analysis of reference materials of the analyte(s) of interest to achieve the required performance of the instrument.
Development and Optimization of Sample Preparation Techniques | 5.3 | Requirement | The sample preparation technique shall be evaluated and optimized using reference materials of the analyte(s) of interest. The primary goal is to demonstrate that the sample preparation steps allow for adequate extraction, detection, identification, and/or quantitation of the analyte(s). Sample preparation shall be evaluated with fortified matrix samples.
Establishing a Validation Plan | 6 | Section Title
Establishing a Validation Plan | 6 | Requirement | A validation plan shall be in place prior to starting any validation experiments. The validation plan is separate from a laboratory’s standard operating procedure for method validation. The plan shall include the instrumental method(s) and sample preparation technique(s) to be used for a specific method. Further, it shall document the validation requirements of the method, as well as the limits of the method that will allow it to be fit for use. The validation plan provides direction for the experiments that will be performed and acceptance criteria for each parameter. Annex A and Annex B provide examples of validation plans.
Required Validation Parameters Based on Scope of the Method | 7 | Section Title
Required Validation Parameters Based on Scope of the Method | 7.1 | Requirement | The scope of forensic toxicology methods is typically categorized as screening methods, qualitative confirmation/identification methods, or quantitative methods. As such the following validation parameters shall be evaluated.
Screening (Immunoassay-based) | 7.2 | Requirement | Screening (Immunoassay-based): a) limit of detection; b) precision (at the decision point); c) processed sample stability (if applicable).
Screening (all other methods targeting specific analytes or analyte classes) | 7.3 | Requirement | Screening (all other methods targeting specific analytes or analyte classes): a) interference studies; b) limit of detection; c) ionization suppression/enhancement [for applicable techniques, such as liquid chromatography/mass spectrometry (LC/MS)]; d) processed sample stability (if applicable).
Qualitative confirmation/identification | 7.4 | Requirement | Qualitative confirmation/identification: a) carryover; b) interference studies; c) ionization suppression/enhancement (for applicable techniques, such as LC/MS); d) limit of detection; e) processed sample stability (if applicable).
Quantitative analysis | 7.5 | Requirement | Quantitative analysis: a) bias; b) calibration model; c) carryover; d) interference studies; e) ionization suppression/enhancement (for applicable techniques, such as LC/MS); f) limit of detection; g) limit of quantitation; h) precision; i) dilution integrity (if applicable); j) processed sample stability (if applicable).
Specific Requirements for Conducting Method Validation Experiments | 8 | Section Title
General | 8.1 | Requirement | All validation experiments shall be conducted using fortified samples for each matrix type which the method is intended, unless otherwise noted. For example, a method with a scope to analyze both postmortem blood and urine specimens shall include complete validation experiments using fortified blank postmortem blood samples and complete validation experiments using fortified blank urine samples. In some instances, it may be more appropriate to analyze previously characterized human samples instead of fortified samples for selected method validation studies (e.g., interference studies). Validation studies shall be conducted in a manner similar to casework. This may include conducting validation studies on different days, by different analysts, on all identical instruments to be utilized for the assay, and ensuring that the instruments meet the same daily performance requirements as for casework. Fortified matrix samples should be prepared from reference materials that are from a different source (e.g., supplier or lot number) than used to prepare calibration samples. In instances where the same source shall be utilized, separate weighings or solutions shall be used to prepare these samples. The following requirements are the minimum for assessing the listed validation parameters in forensic toxicology methods. They are listed alphabetically and not necessarily in procedural order. Some of the validation experiments are demonstrated in Annex A and Annex B. Annexes C, D, and E provide guidance on how to efficiently perform validation experiments.
Bias and Precision1 | 8.2 | Section Title | 1. An accurate measurement is one with acceptable bias and precision.
Bias | 8.2.1 | Requirement | Bias studies shall be carried out for all quantitative methods. These can be conducted concurrently with precision studies. Bias shall be measured in pooled fortified matrix samples using a minimum of three separate samples per concentration at three different concentration pools (low, medium and high2) over five different runs3. 2. For purposes of this document, low concentrations shall be no more than approximately 3 times the lowest end of the working range of the method and high concentrations shall be within approximately 80% (or more) of the highest end of the working range of the method, unless otherwise noted. Medium concentrations shall be near the midpoint of the low and high concentrations. 3. In some instances, analyte instability may preclude the ability to use concentration pools of fortified samples (e.g., cocaine in unpreserved whole blood). In these instances, a laboratory may resort to fortifying different samples with each run. The bias shall be calculated for each concentration using the following formula: Bias (%) at Concentrationx = [Grand Mean of Calculated Concentrationx - Nominal Concentrationx/Nominal Concentrationx] x 100 The maximum acceptable bias shall be ±20% at each concentration. For some analyses where less bias is required (e.g., ethanol), a bias of ±10% or better should be expected. It is recommended that the same data used in bias studies also be used for precision calculations.
Precision | 8.2.2 | Section Title
General | 8.2.2.1 | Requirement | Precision studies shall be carried out for all quantitative methods, as well as at the decision point for immunoassays. These studies can be carried out concurrently with bias studies, if required in the validation plan. Precision is expressed as the coefficient of variation (% CV). The mean and standard deviation (std dev) of the response is calculated for each concentration to determine the % CV. %CV = std dev/mean response x 100
Precision of Immunoassays at Decision Point | 8.2.2.2 | Requirement | At a minimum, precision at the immunoassay’s decision point (i.e. cutoff concentration) shall be assessed for the target analyte using three separate samples from three different concentration pools over five different runs: a) generally no more than 50% below the decision point, b) at the decision point, and c) generally no more than 100% above the decision point. The intent of this requirement is to ensure that there is evidence that the assay separates presumptive positive and negative samples at the decision point. For Enzyme Linked Immunosorbent Assays (ELISA), the difference between the absorbance of a negative sample (B0) and the absorbance of a specimen (B) should be used as a percentage: [B/B0] x 100 and not an absolute value. Note that B0 is established for each run. For liquid reagent assays (e.g., enzyme multiplied immunoassay technique [EMIT], cloned enzyme donor immunoassay [CEDIA®4], etc.), the absorbance value may be directly used. a) Run each concentration 3 times in five separate runs. Additionally, for ELISA assays, at least one negative sample per run shall be analyzed to establish B0. b) Calculate the grand mean (n=15) and related grand standard deviation for each concentration pool. c) The %CV shall not exceed 20% at each concentration using all 15 sample results per concentration. d) The grand mean plus or minus two standard deviations of the low and high concentration pools shall not overlap with the mean of the decision point.
Precision of Quantitative Procedures | 8.2.2.3 | Section Title
General | 8.2.2.3.1 | Requirement | For quantitative procedures, two different types of precision studies shall be assessed during method validation: within-run precision and between-run precision. At a minimum, precision shall be assessed using triplicate analyses per concentration pool (low, medium and high) over five different runs. The different runs used to evaluate precision may be performed within the same day, provided a different calibration curve is used for each run. The % CV shall not exceed 20% at each concentration. It is noted that certain analytical methods (e.g., blood alcohol analysis) should require a much lower coefficient of variation (≤10% CV). The largest calculated within-run and between-run % CV for each concentration shall be used to assess precision acceptability.
Within-Run Precision Calculations | 8.2.2.3.2 | Requirement | Within-run precisions are calculated for each concentration separately for each of the five runs. Within-run precision may be calculated using the data from each run’s triplicate analyses at each concentration as: Within-run CV(%) = std dev of a single run of samples/mean calculated value of a single run of samples x 100
Between-Run Precision Calculations | 8.2.2.3.3 | Requirement | Between-run precision is calculated for each concentration over the five runs. This may be done by using the combined data from all replicates of each concentration as: Between-Run CV(%) = std dev of all observations for each concentration/grand mean for each concentration x 100
One-Way Analysis of Variance (ANOVA) Approach to Calculate Combined Within- Run and Between-Run Precision | 8.2.2.3.4 | Informational Text | Both within-run and between-run precisions may be calculated using the one-way ANOVA approach with the varied factor (run number) as the grouping variable. The ANOVA calculations can be easily performed using a spreadsheet or a statistical software program. Using this approach, within-run precisions are calculated for each concentration as: Within-run CV(%) = [square root MSwg/grand mean for each concentration] x 100 where MSwg is the mean square within groups obtained from the ANOVA table. Likewise, between-run precisions are calculated as: Between-Run CV(%) = [(square root MSbg+(n-1)*MSwg/n)/grand mean for each concentration] x 100 where MSbg is the mean square between groups obtained from the ANOVA table and n is the number of observations in each group (e.g., n=3 if doing triplicate analyses). Annex A provides an example of how the ANOVA approach may be used to calculate within-run and between-run precision.
Calibration Model | 8.3 | Requirement | The calibration model shall be determined for all quantitative methods. This is accomplished by first determining the working range. Within this range, there will be a correlation between signal response (e.g., peak area ratio of analyte and internal standard) and analyte concentration in the sample. The calibration model is the mathematical model that describes this correlation. The choice of an appropriate model (i.e., linear or quadratic) is necessary for accurate and reliable quantitative results. Calibrator samples are analyzed to establish the calibration model. The use of matrix-matched calibrator samples is encouraged, but not required. A single source of blank matrix (per matrix type) may be used when experimentally establishing the appropriate calibration model for a method. Regardless of the matrix used to prepare calibrator samples, a laboratory shall demonstrate acceptable bias and precision with control samples prepared in all matrices intended to be analyzed by the method (see Section 7.1). For example, blood alcohol methods may demonstrate acceptable bias and precision in whole blood controls using aqueous calibrator samples. Likewise, acceptable bias and precision may be demonstrated using calibrator samples prepared in whole blood but used to quantitate analytes in different matrices (e.g., postmortem tissues, serum, urine). The calibrator samples shall span the range of concentrations expected in day-to-day operations. At least six different non-zero concentrations shall be used to evaluate the calibration model.
Calibration Model | 8.3 cont'd 1 | Requirement cont'd | The concentrations shall be appropriately spaced across the calibration range to establish the most appropriate calibration model. A minimum of five replicates per concentration shall be used. The replicates to establish the calibration model may be in the same or in separate runs. All data points from the five replicates shall be plotted together (using a statistical software package) to establish the calibration model. The origin shall not be included as a calibration point. The simple linear regression model using the least squares method is the most often used calibration model. However, this model is only applicable when there is constant variance over the entire concentration range (homoscedasticity). When there is a notable difference between variances at the lowest and highest concentrations, a weighted least squares model or other appropriate non-linear model should be applied5. 5 In general, non-linear models may require additional calibrators to accurately characterize the curve. This is generally the case when the concentration range exceeds one order of magnitude. Ultimately, the simplest calibration model that best fits the concentration-response relationship should be used. A calibration model shall not be evaluated simply via its correlation coefficient (r).
Calibration Model | 8.3 cont'd 2 | Requirement cont'd | Instead, a calibration model should be visually evaluated using residual plots. These allow one to check for outliers that may be eliminated if found to be statistically significant (e.g., outside ±3 standard deviations). Further, residual plots allow one to determine if the variances appear to be equal across the calibration range with a similar degree of scatter at each concentration. They also give an indication if the chosen model adequately fits the data. For example, random distribution of individual residuals around the zero line suggests that a linear model is appropriate. Finally, there are other appropriate alternatives to evaluate calibration models (i.e., ANOVA lack-of-fit test for unweighted linear models, checking for significance of the second order term in quadratic models, assessment of coefficient of determination (r2) for linear models). If a linear calibration model has been established, fewer calibration samples (i.e., fewer levels or single/fewer replicates) may be used for routine analysis. However, if fewer calibration samples are chosen, the same calibrators (e.g., number, replicates, and concentration levels) shall be used to construct the calibration curves for the bias and precision studies. Further the calibration data shall include the lowest and highest calibration levels used to establish the model, as well as include no fewer than four non-zero calibration points. Additionally, once the calibration model is established for a validated method, it shall not be arbitrarily changed to achieve acceptable quality control results during a given analytical run. For example, one shall not switch from an unweighted linear model to a weighted linear model in order to adjust for changes in instrument performance without revalidation of the impact on the change.
Carryover | 8.4 | Requirement | Analyte carryover into a subsequent sample may lead to an inaccurate qualitative or quantitative result when using instrumental methods. Carryover shall be evaluated during method validation intended for confirmation and/or quantitation, unless a laboratory is constantly addressing carryover in their quality assurance (QA)/quality control (QC) practices. To evaluate carryover as part of method validation, blank matrix samples are analyzed immediately after a high concentration sample or reference material. The highest fortified concentration at which no analyte carryover is observed (above the method's LOD) in the blank matrix sample is determined to be the concentration at which the method is free from carryover. This carryover concentration for each analyte in the method shall be confirmed using triplicate analyses. It is acceptable to limit the carryover study to the highest point of your calibration curve in quantitative assays. The analytical procedure should be modified to remove any carryover. In cases when it is not possible to eliminate the carryover, the SOP shall address how carryover will be managed.
Interference Studies | 8.5 | Section Title
General | 8.5.1 | Requirement | Interfering substances from common sources shall be evaluated in all screening (except immunoassays), qualitative identification, and quantitative methods.
Evaluating Matrix Interferences | 8.5.2 | Requirement | Whenever possible, blank matrix samples from a minimum of ten different sources without the addition of an internal standard (when used in the method) shall be analyzed to demonstrate the absence of common interferences from the matrix. While this approach may detect the more common matrix interferences, it is recognized that less common interferences may not be detected.
Evaluating Interferences from Stable-Isotope Internal Standards | 8.5.3 | Requirement | For methods employing stable isotope internal standards, the isotopically-labeled compounds may contain the non-labeled compound as an impurity. Additionally, the mass spectra of the labeled analogs may contain fragment ions with the same mass-to-charge ratios as the significant ions of the target analyte. In both instances, analyte identification or quantitation could be impacted. Stable-isotope internal standard interferences shall be assessed by analyzing a single blank matrix sample (per matrix type) fortified with the internal standard(s) and monitoring the signal of the analyte(s) of interest. Interferences below the LOD of the assay may be insignificant depending on the laboratory’s mission. Likewise, a single blank matrix sample (per matrix type) fortified with the analyte(s) at a concentration near the upper limit of the calibration range shall be analyzed without internal standard to evaluate whether relevant amounts of the unlabeled analyte ions appear as isotopically-labeled compound fragments which could impact quantitation.
Evaluating Interferences from Other Commonly Encountered Analytes | 8.5.4 | Requirement | For all methods other than immunoassays, it is necessary to evaluate other analytes which may be expected to be present in case samples for their potential to interfere with the method’s analytes. For example, a method developed to analyze blood for cocaine shall evaluate whether other common drugs of abuse, metabolites, and structurally-similar compounds interfere with the assay. Likewise, a headspace gas chromatograph-flame ionization detection (GC-FID) method developed for ethanol shall evaluate whether other common volatile organic compounds interfere with the assay. This evaluation shall be accomplished by analyzing fortified matrix samples, previously analyzed case samples, or neat reference materials of the potential interference(s) at high therapeutic or lethal concentrations, depending on the analyte, the matrix, and the laboratory’s mission. The most common drugs/metabolites encountered in the laboratory shall be included in the evaluation together with other common drugs within the classification, where appropriate.
Ionization Suppression/Enhancement | 8.6 | Section Title
General | 8.6.1 | Requirement | The enhancement or suppression of analyte ionization resulting from the presence of co-eluting compounds is a phenomenon commonly encountered in LC/MS applications. When average suppression/enhancement of the analyte’s target ion (or ion transition and qualifying ions) exceeds ±25% or the % CV of the suppression/enhancement exceeds 20%, a laboratory shall demonstrate that there is no impact on other critical validation parameters. For example, suppression/enhancement of ionization is most likely to impact the LOD of a qualitative method. Likewise, the LOD and LLOQ may be affected by ionization suppression/enhancement in quantitative methods. The influence on the above parameters shall be assessed by at least tripling the number of different sources of blank matrices used in their evaluation. For example, if the average suppression/enhancement exceeds ±25%, the LOD determination shall be performed in at least nine unique sources of blank matrices. Laboratories shall also assess the impact of ionization suppression/enhancement on the method’s internal standards. Ionization suppression/enhancement shall be evaluated using either of the approaches that follow.
Post-column Infusion to Assess Ionization Suppression/Enhancement | 8.6.2 | Requirement | This approach provides information on retention times where ionization suppression/enhancement occurs. It is useful during method development, but may also be used to assess the amount of ionization suppression/enhancement for LC/MS based confirmation methods. Solutions at both low and high concentrations of the analyte are individually infused with a syringe pump into the eluent from the column via a post-column “T”-connector and a constant baseline signal for the analyte of interest is monitored. Whenever possible, a minimum of ten different processed blank matrix samples (per matrix type) that are representative of the quality of samples typically encountered in casework are injected into the LC/MS during infusion of the solutions.6 6 Additional matrix samples may be needed in postmortem toxicology given the variety of sample conditions typically encountered in this work. If there is any considerable suppression/enhancement (>25%) of the infused analyte signal at the retention time of the analyte, then modification of the chromatographic system or the sample preparation may be required to minimize the effect of ionization suppression/enhancement.
Post-Extraction Addition to Assess Ionization Suppression/Enhancement | 8.6.3 | Requirement | This approach yields a quantitative estimation of ionization suppression/enhancement. It is useful for assessing the amount of ionization suppression/enhancement for LC/MS based methods. Two different sets of samples shall be prepared and the analyte peak areas of neat standards shall be compared to matrix samples fortified with neat standards after extraction or processing. Set one consists of neat standards prepared at two concentrations – one low and one high. Each of these neat standards shall be injected a minimum of six times to establish a mean peak area for each concentration. Set two should consist of a minimum of ten different matrix sources (per matrix type).7 7 Additional matrix samples may be needed in postmortem toxicology given the variety of sample conditions typically encountered in this work. Each matrix source shall be extracted in duplicate. After the extraction is complete, each matrix sample shall be reconstituted/fortified with either the low or high concentration neat standard. The average area of each set ( ̅X) shall be used to estimate the suppression/ enhancement effect at each concentration as follows: Ionization suppression or enhancement (%) = (( ̅X Area of Set 2/ ̅X Area of Set 1) - 1) x 100 Two ionization suppression/enhancement percentages shall be established – one at the low concentration and one at the high concentration.
Limit of Detection | 8.7 | Section Title
General | 8.7.1 | Requirement | LOD studies shall be carried out for all methods. There are a number of different approaches for determining the LOD. Select the approach that provides the most reasonable estimation of the LOD given the analytical instrumentation (or lack thereof) utilized in the method. A method's LOD incorporates instrumental performance, as well as the sample matrix and inherent procedural limitations. Therefore, the LOD shall be assessed over multiple runs using fortified matrix samples from at least three different sources of blank matrix, unless otherwise indicated. For example, if the assay is to be used for postmortem blood samples, three independent representative postmortem blood sources are needed. The LOD shall be determined by one of the following approaches.
Estimating LOD for Immunoassays | 8.7.2 | Requirement | The laboratory may assign the decision point (i.e., cutoff concentration) as the LOD for immunoassays. Most of these assays are known to cross-react with numerous drugs (e.g., benzodiazepines, opiates, amphetamines) and metabolites. When a laboratory declares to their customers that they are able to detect specific analytes demonstrating low cross-reactivity (less than or equal to the target analyte) using the immunoassay, they shall verify their ability to reliably detect these compounds. As an example, a benzodiazepine immunoassay targeted for oxazepam typically has low cross-reactivities to many other benzodiazepines. If a laboratory uses the decision point (i.e. cut-off concentration) determined by the manufacturer, the laboratory shall mathematically estimate the LOD concentration for any benzodiazepines that cross-react at less than 100% and that they declare to their customers they are able to detect with the immunoassay. For example, if lorazepam cross-reacts at 50%, it would have an equivalent decision point concentration of 100 ng/mL in an oxazepam-targeted assay with a decision point of 50 ng/mL. In contrast, if another benzodiazepine (e.g., alprazolam) has greater cross-reactivity than the target drug, oxazepam, there is no requirement to mathematically estimate the LOD concentration for that drug provided that the decision point for the drug (alprazolam) is not lower than the decision point (50 ng/mL) for the target analyte (oxazepam). In contrast, if a laboratory uses a different decision point than that recommended by the manufacturer of an immunoassay kit, the laboratory shall experimentally determine the LOD of the other benzodiazepines in the panel with cross-reactivities at or below that of the target analyte. This evaluation may require an adjustment or reevaluation of the decision point or the target compound depending on the needs and mission of the laboratory. The experiments performed to assess precision (Section 8.2.2.2) shall be used to demonstrate a suitable LOD has been determined.
Estimating LOD for Immunoassays | 8.7.2 cont'd 1 | Requirement cont'd | The replicates for analytes with cross-reactivities at or below the target analyte (n=15) must show an average that indicates a “positive” response compared to the target analyte. A single source of blank matrix (per matrix type) may be used to estimate the LOD of immunoassays.
Estimating LOD for a Non-Instrumental Method | 8.7.3 | Requirement | This approach is most often used when screening for the presence or absence of a specified analyte or class of analytes (e.g., color tests). To estimate the LOD for a non-instrumental method, at least three different blank matrix sources fortified with decreasing concentrations of analyte shall be analyzed over a minimum of three runs. Multiple analysts should be involved in estimating the LOD using this approach, due to greater subjectivity in these methods. The lowest concentration of analyte that yields a positive result on all runs and observed by all participating analysts shall be considered the LOD.
Using the Lowest Non-Zero Calibrator as the LOD | 8.7.4 | Requirement | This technique is useful for quantitative methods. In some instances, it may be sufficient to define the LOD as the value of the lowest non-zero calibrator. For each matrix type, at least three different blank matrix sources shall be fortified with the analyte at the concentration of the lowest calibrator and analyzed over a minimum of three runs to demonstrate that all detection and identification criteria are met. If desired, it is acceptable to use the same calibrator replicates used to establish the calibration model (Section 8.3) for this approach, but additional samples/replicates may be needed to meet the minimum of nine data points including at least three sources per matrix type.
Using the Decision Point Concentration as the LOD | 8.7.5 | Requirement | This technique is useful for qualitative and quantitative methods. In some instances, it may be sufficient to define the LOD as the value of an administratively-defined decision point. For example, a laboratory may choose to define a method’s LOD for ethanol as 0.02 g/dL for blood based on the laboratory’s administratively defined decision point for reporting this analyte, even though a lower LOD is analytically achievable. For each matrix type, at least three different blank matrix sources shall be fortified with the analyte at the concentration of the decision point and analyzed over a minimum of three runs to demonstrate that all detection and identification criteria are met.
Estimating LOD Using Background Noise | 8.7.6 | Section Title
General | 8.7.6.1 | Requirement | These approaches for determining LOD are only useful for instrumental methods that demonstrate background noise. A minimum of three different blank sources per matrix type shall be used.
Estimating LOD Using Reference Materials | 8.7.6.2 | Requirement | Three (or more) sources of blank matrix samples (per matrix type) fortified at decreasing concentrations shall be analyzed in duplicate (two separate replicates) for at least three runs. The LOD shall be the lowest concentration that 1) yields a reproducible instrument response greater than or equal to 3.3 times8 the noise level of the background signal in an area around the analyte peak and 2) achieves acceptable predefined detection and identification criteria (e.g., retention time, peak shape, mass spectral ion ratios) for all replicates at that concentration. 8 Use of 3.3 in the LOD calculation provides a false positive error rate of 0.0005%. (see Boyd, RK, Basic C, and Bethem RA. Trace Quantitative Analysis by Mass Spectrometry. Hoboken, N.J.: John Wiley (2008). While it may be possible to visually assess the signal-to-noise ratio, such an approach is subjective. Therefore, the signal-to-noise ratio shall be manually calculated or determined through use of instrumentation software. If manually calculated, the signal is defined as the height response of the analyte peak and the noise is defined as the amplitude between the highest and lowest point of the baseline in an area around the analyte peak. Each replicate shall be independently evaluated. Signal-to-Noise = height of analyte/amplitude of noise
Estimating LOD Using Statistical Analysis of Background | 8.7.6.3 | Requirement | To determine the LOD using this approach, a minimum of three sources of blank matrix shall be analyzed in duplicate (two separate samples) over at least three runs. The average signal (e.g., integrated area of signal at the analyte’s retention time) from all blank matrix samples and the standard deviation of the signal shall be calculated. Likewise, fortified matrix samples of decreasing concentration shall be analyzed in duplicate over the course of at least three runs. The lowest concentration of a fortified matrix sample that consistently yields a signal greater than the average signal from the blank matrix samples plus 3.3 times the standard deviation of the signal from the blank matrix samples, shall be identified as the LOD.
Estimating LOD Using a Linear Calibration Curve | 8.7.7 | Requirement | This technique is useful for any quantitative method that follows a linear calibration model. A minimum of three independent calibration curves shall be constructed across the working range of the analytical method over different runs. A single source of blank matrix (per matrix type) may be used to estimate the LOD using this approach. The LOD shall be estimated from the standard deviation of the y intercept (sy) and the average slope (Avgm) as: LOD = (3.3 sy)/Avgm
Lower Limit of Quantitation | 8.8 | Section Title
General | 8.8.1 | Requirement | LLOQ studies shall be carried out for all quantitative methods. There are a number of different approaches for determining a method's LLOQ. Select the approach that provides the most reasonable estimation of the LLOQ given the analytical instrumentation utilized in the method. A method's LLOQ incorporates instrumental performance, as well as the sample matrix and inherent procedural limitations. Therefore, the LLOQ shall be assessed over multiple runs using fortified, matrix samples from at least three different sources per matrix type.
Using the Lowest Non-Zero Calibrator as the LLOQ | 8.8.2 | Requirement | In some instances, it may be sufficient to define the LLOQ as the value of the lowest non-zero calibrator. For each matrix type, at least three different blank matrix sources shall be fortified with the analyte at the concentration of the lowest calibrator and analyzed over a minimum of three runs to demonstrate that all detection, identification, bias, and precision criteria are met. If desired, it is acceptable to use the same calibrator replicates used to establish the calibration model (Section 8.3) for this approach, but additional samples/replicates may be needed to meet the minimum of nine data points including three sources per matrix type.
Using Decision Point Concentration as the LLOQ | 8.8.3 | Requirement | In some instances, it may be sufficient to define the LLOQ as the value of an administratively defined decision point. For example, a laboratory may choose to define a method’s LLOQ for GHB as 5 mg/L for antemortem blood based on the laboratory’s administratively defined decision point for reporting this analyte, even though a lower LLOQ is analytically achievable. The concentrations used for this approach shall remain within the previously established calibration curve. For each matrix type, at least three different blank matrix sources shall be fortified with the analyte at the concentration of the decision point and analyzed over a minimum of three runs to demonstrate that all detection, identification, bias, and precision criteria are met.
Additional Validation Parameters | 9 | Section Title
General | 9.1 | Requirement | In certain instances, it is important to evaluate additional validation parameters, if they are applicable. These include processed sample stability of the analyte(s) and the effect of sample dilution on bias and precision. A laboratory shall include these parameters in their validation plan, and determine if they are applicable to the analytical method or if they are already addressed through other means (e.g.., quality assurance practices). The laboratory validation plan shall include documentation of this evaluation.
Dilution Integrity | 9.2 | Requirement | The effect of sample dilution shall be determined during validation of quantitative methods, if this is a routine practice within the laboratory. At times, this may be due to low specimen volume requiring the sample or assay to be adjusted appropriately. In other instances, excessively high concentrations may be encountered that are above the established calibration range. To bring the analyte concentration within the validated concentration range, the laboratory procedure may allow for reanalysis after dilution of the sample. If dilution of a sample is allowed because of high analyte concentration or low sample volume, then the laboratory shall evaluate the effect of dilution on the method's bias and precision with at least one concentration pool. This shall be accomplished by establishing bias and within-run precision studies (as described in Section 8.2.) at a common dilution ratio (e.g., 1:2) utilized by the laboratory and determining if performance criteria are still met.
Processed Sample Stability | 9.3 | Requirement | Processed samples are typically analyzed in batches; however, circumstances may arise in which they cannot be analyzed within a reasonable amount of time due to atypical events (e.g., instrument failures or loss of power). It may be necessary to analyze processed samples the following day or even later. In these instances, the laboratory shall evaluate the length of time processed samples can be maintained before they undergo unacceptable changes. The following approach provides a means of evaluating loss of analytes in stored, processed samples at both low and high concentrations that could impact the ability to accurately detect, identify, and quantitate them. Processed sample stability experiments shall utilize sets of blank matrix samples fortified at both low and high concentrations. A single source of blank matrix (per matrix type) may be used to evaluate processed sample stability. The samples may be prepared by the laboratory, purchased from a commercial source, or appropriately pooled patient samples. A large enough volume of each of these samples should be used in order to complete the studies. Numerous aliquots from each concentration set shall be processed (i.e., extracted) using the method under validation. The processed samples for a given concentration pool shall be combined, mixed well, and then divided into different (autosampler) vials for instrumental analysis. The first vials of each concentration shall be immediately analyzed in triplicate to establish the time zero responses. All remaining vials shall be maintained in a manner that they would typically be stored during routine analysis (e.g., at refrigerated or room temperature on autosampler).
Processed Sample Stability | 9.3 cont'd 1 | Requirement cont'd | The remaining vials shall be analyzed in triplicate at different time intervals that represent the typical time range expected for processed samples to wait before being injected into the instrument. The analyte shall be considered stable until the average signal (e.g., peak area or ratios of peak area of analyte to internal standard) compared to the time zero average signal falls outside of the method’s acceptable bias. As an example, say that a method’s bias limit is ±15% and the time zero average signal is 100,000. The laboratory’s processed samples are placed into different autosampler vials and are analyzed repeatedly up to 72 hours. For this example, the processed sample’s analyte is considered stable until the average signal falls outside of the 85,000 – 115,000 range (±15% of the time zero average signal).A plot of the average response for each concentration pool against each time point with linear regression allows for an assessment of trends. Addendum B provides a detailed example of this approach.
Required Revalidation of Previously Validated Methods | 10 | Section Title
Required Revalidation of Previously Validated Methods | 10 | Requirement | Modifications to a validated method shall be evaluated to confirm that the changes do not have an adverse effect on the method’s performance. The decision regarding which performance characteristics require additional validation shall be based on consideration of the specific parameters likely to be affected by the change(s). These changes may include, but are not limited to: a) analytical conditions, b) instrumentation, c) sample processing, d) data software. For example, changes of extraction solvent or buffer may affect linearity, interferences, LLOQ, precision, and bias. A change of the analytical column stationary phase or a change in mobile phase composition may affect linearity and interferences. Further, consideration should be given to conducting parallel studies with known or proficiency samples utilizing both a previously validated method and the modified method to evaluate the effects of the changes. The goal should be to demonstrate the impact the changes have on the performance of the previously validated procedure.
Documentation Requirements for Method Validation | 11 | Section Title
Documentation Requirements for Method Validation | 11 | Requirement | Record keeping is an essential part of laboratory operating procedures and is a key component of method validation. The data generated during method validation studies shall be maintained and available for audits, reviews, or inspections. These records shall be organized for easy retrieval and review. Method validation records shall include a summary of the validation studies conducted and their results. The format of this summary report may be a brief bulleted report or in a table to facilitate a swift review of validation studies. The summary shall minimally include the following: a) scope; b) validation plan; c) description of all the parameters evaluated, if any of the parameters were not evaluated, then the reason shall be stated or justified; d) sample preparation steps to include concentrations and matrices; e) raw data or reference to where the raw data are stored; f) results and calculations; g) conclusions; h) references; i) documentation of management review and approval. The validation records shall also contain specific details regarding the studies conducted, including: a) individuals involved in the method validation, b) specific instrumentation, c) dates.
Documentation Requirements for Method Validation | 11 cont'd 1 | Requirement cont'd | Method validation documentation shall also include a copy of the newly developed analytical method or a reference to its location. Further, validation documentation should be retained for a minimum of 10 years after the method is retired.
Efficiency with Validation | 12 | Section Title
Efficiency with Validation | 12 | Informational Text | Keep in mind that some validation experiments may be conducted concurrently with the same fortified samples. Annex C, D, and E present example approaches to assist in streamlining validation experiments.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics