## Instructions for Use
AAFS Standards Resources and Training
This checklist is made possible through the following financial assistance award 70NANB24H078 awarded to AAFS from U.S. Department of Commerce, National Institute of Standards and Technology Each checklist is specific to a version of a standard, which is hyperlinked in the checklist. Not all sections of a standard are included in a checklist. A checklist is a supplementary tool and should not be used independently. To properly use a checklist, reference the complete standard. Go to https://www.nist.gov/organization-scientific-area-committees-forensic-science/access-standards for information on how to access published standards. A checklist does not provide interpretation guidance. If you have questions related to interpretation, please contact the standard developing organization (SDO). The SDO may have standard specific training available.
A checklist can be used for one or both of the following purposes: • to evaluate the implementation of a standard, or • to audit (internal or external) conformance with requirements in the standard and an FSSP’s management system
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions and an on-demand training webinar.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASB 056-2025 1st Ed
ANSI/ASB 056-2025, First Edition
Standard for Evaluation of Measurement Uncertainty in Forensic Toxicology
Discipline: | Toxicology
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Requirements for Measurement Uncertainty for Quantitative Determinations | 5 | Section Title
General Requirements | 5.1 | Section Title
General Requirements | 5.1.1 | Requirement | Testing laboratories and breath alcohol programs shall have and apply procedures for evaluating MU for test methods that produce a quantitative test result and for methods used to calibrate breath alcohol measuring instruments.
General Requirements | 5.1.2 | Requirement | Records of MU evaluations shall be maintained.
General Requirements | 5.1.3 | Requirement | MU shall be evaluated for each measurement process and is specific to the measurement process. This includes, but is not limited to:
General Requirements | 5.1.3.1 | Requirement | Each calibration method shall be evaluated separately.
General Requirements | 5.1.3.2 | Requirement | Each combination of analyte, extraction, and analytical technique shall be evaluated separately. NOTE 1 MU specific to each measurement process means not using the largest evaluated MU for more than one analyte within a method or one analyte across methods. NOTE 2 Statistical data evaluation may indicate a need to evaluate different matrices separately.
General Requirements | 5.1.4 | Requirement | Test and calibration methods for which the MU is evaluated shall meet the minimum requirements set forth in: a) ANSI/ASB Standard 017, Standard for Metrological Traceability in Forensic Toxicology. b) ANSI/ASB Standard 036, Standard Practices for Method Validation in Forensic Toxicology. c) ANSI/ASB Standard 054, Standard for a Quality Control Program in Forensic Toxicology Laboratories. d) ANSI/ASB Standard 055, Standard for Breath Alcohol Measuring Instrument Calibration.
Step 1—Specify the Measurement Process | 5.2 | Requirement | The measurand shall be defined. NOTE This can be a written statement, a visual diagram, and/or a mathematical expression. Be specific when defining the measurand. EXAMPLES Testing of biological samples Concentration of ethanol (g/100mL) in antemortem whole blood using GC-FID Concentration of oxycodone (mg/kg) in a sample of liver homogenate using LC-MS/MS Calibration of breath alcohol measuring instruments Calibration of XYZ model breath alcohol measuring instrument using dry gas certified reference material.
Step 2—Identify Uncertainty Components | 5.3 | Requirement | Minimum method components that shall be considered, as applicable, in an evaluation of MU include: a) certified reference material(s) and calibrations of equipment used to establish metrological traceability; b) data from the measurement process (i.e. , repeatability, reproducibility or from intermediate measurement conditions); c) human factors (e. g. , multiple personnel performing the same measurement method, experience, training); d) sampling conducted during the measurement method; e) sample preparation; and f) environmental conditions during the measurement process.
Step 3—Quantify Uncertainty Components | 5.4 | Section Title
General | 5.4.1 | Section Title
General | 5.4.1.1 | Requirement | Uncertainty components shall be quantified.
General | 5.4.1.2 | Requirement | No fewer than three significant figures shall be carried through all calculations to ensure reporting requirements can be met.
General | 5.4.1.3 | Requirement | The method of evaluation, Type A or Type B, shall be determined for each component. NOTE 1 It is most common to use a mixture of the two methods, where some uncertainty components are quantified using a Type A method of evaluation and some uncertainty components are quantified using a Type B method of evaluation. NOTE 2 Double counting of a component will result in overestimating the measurement uncertainty.
Minimum Requirement(s) for Type A Evaluations | 5.4.2 | Section Title
General | 5.4.2.1 | Requirement | Testing laboratories and breath alcohol programs shall specify in their procedure the source(s) of the Type A data to be used.
Testing Laboratories | 5.4.2.2 | Section Title
Selection of Type A Data | 5.4.2.2.1 | Section Title
Selection of Type A Data | 5.4.2.2.1.1 | Informational Text | Validation data may initially be used to evaluate one or more specific Type A uncertainty components.
Selection of Type A Data | 5.4.2.2.1.2 | Requirement | Control data shall be used for the Type A uncertainty component after validation and implementation.
Selection of Type A Data | 5.4.2.2.1.3 | Requirement | Proficiency test data may also be used for a Type A uncertainty component; however, if used, the test(s) shall have established metrological traceability. NOTE A consensus result does not establish metrological traceability.
Selection of Type A Data | 5.4.2.2.1.4 | Requirement | Data used in Type A evaluations shall: a) be representative of the measurand that will be tested; b) be representative of the range (e.g., matrix, detector response over the expected concentration range) of the measurements made; c) be representative of the data generated during ongoing analysis by personnel who have demonstrated competence; and d) be evaluated according to the size and distribution of the statistical sample. NOTE Approaches to selecting Type A Data include, but are not limited to: ⎯ using control data generated since method implementation; ⎯ using a laboratory-specified number of control data points from the most recent analyses; or ⎯ using control data from only the current analytical batch in non-routine analyses with limited data points.
Calculation of the Quantity Value for Type A Data | 5.4.2.2.2 | Section Title
Calculation of the Quantity Value for Type A Data | 5.4.2.2.2.1 | Requirement | The standard deviation or relative standard deviation shall be calculated using data for each Type A uncertainty component. NOTE 1 Method performance is typically represented by measurements of control samples taken over multiple batches, each with different calibrations. NOTE 2 If multiple replicates of a control level are available per batch, the data from all replicates may be included when calculating the standard deviation or relative standard deviation. Including all data in the standard deviation calculation will bias the standard uncertainty slightly if the data exhibits any batch-to-batch variation but mitigates the need for more complex standard deviation calculations. This would provide an assessment of the Type A uncertainty that is either on target or conservative (i. e. , overestimated) for the reported specimen value. NOTE 3 If needed, other statistical methods, such as the ANOVA method or random subsampling of the data to select a single instance from each batch, can be used to correct this bias.
Calculation of the Quantity Value for Type A Data | 5.4.2.2.2.1.1 | Requirement | When the result to be reported for a specimen is either an individual measured value or the average of multiple measured values from a single instrumental batch, the standard deviation or the relative standard deviation shall be used as the Type A standard uncertainty for the reported specimen value.
Calculation of the Quantity Value for Type A Data | 5.4.2.2.2.1.2 | Requirement | When the result to be reported for a specimen is the average of measured values from multiple instrumental batches, the standard deviation or the relative standard deviation divided by the square root of the number of instrumental batches used when averaging the specimen data shall be used (i.e. , standard deviation of the mean of multiple batches) as the Type A standard uncertainty for the reported specimen values.
Calculation of the Quantity Value for Type A Data | 5.4.2.2.2.2 | Requirement | Multiple controls within the same method. Testing laboratories shall evaluate variance of control data (e. g. , perform a statistical F-test).
Calculation of the Quantity Value for Type A Data | 5.4.2.2.2.2.1 | Requirement | If consistent variance is demonstrated, testing laboratories shall: a) combine data from all controls analyzed; or b) select data from one specified control (e. g. , a control at or near a legal specification).
Calculation of the Quantity Value for Type A Data | 5.4.2.2.2.2.2 | Requirement | If consistent variance is not demonstrated, testing laboratories shall: a) utilize the Type A data from the control producing the largest variance; or b) perform an in-depth evaluation to determine where the variation change occurs and establish an appropriate uncertainty to report for each range.
Calculation of the Quantity Value for Type A Data | 5.4.2.2.2.3 | Requirement | Multiple instruments and/or laboratories. To calculate a single MU by combining data from multiple instruments and/or in multiple laboratories, control acceptance and reporting criteria shall be the same across all instruments and laboratories.
Calibration of Breath Alcohol Measuring Instruments | 5.4.2.3 | Section Title
Selection of Type A Data | 5.4.2.3.1 | Section Title
Selection of Type A Data | 5.4.2.3.1.1 | Informational Text | Validation data may initially be used to evaluate one or more specific Type A uncertainty components.
Selection of Type A Data | 5.4.2.3.1.2 | Requirement | Reference material data generated during calibrations shall be used for the Type A uncertainty component after validation and implementation. Reference material data generated during control testing may be used in addition to that generated during calibrations.
Selection of Type A Data | 5.4.2.3.1.3 | Requirement | Proficiency test data may also be used for a Type A uncertainty component; however, the test shall have established metrological traceability if used. NOTE A consensus result does not establish metrological traceability.
Selection of Type A Data | 5.4.2.3.1.4 | Requirement | Data used in Type A evaluations shall: a) be representative of the measurand that will be calibrated; b) be representative of the range of the measurements made; c) be representative of the data generated during ongoing calibrations performed by personnel who have demonstrated competence; and d) be evaluated according to the size and distribution of the statistical sample. NOTE Approaches to selecting Type A data include, but are not limited to: ⎯ using reference material data generated since method implementation; ⎯ using a breath alcohol program specified number of reference material data points from the most recent calibrations; or ⎯ using reference material data from only the current calibration.
Calculation of the Quantity Value for Type A Data | 5.4.2.3.2 | Section Title
Calculation of the Quantity Value for Type A Data | 5.4.2.3.2.1 | Requirement | The standard deviation or relative standard deviation shall be calculated using data for each identified Type A uncertainty component.
Calculation of the Quantity Value for Type A Data | 5.4.2.3.2.2 | Requirement | Calibration method (multiple measurement standards). Breath alcohol programs shall evaluate the variance between measurement standard data by performing a statistical test (e.g. , perform a statistical F-test).
Calculation of the Quantity Value for Type A Data | 5.4.2.3.2.2.1 | Requirement | If a consistent variance is demonstrated, breath alcohol programs shall: a) combine data from all measurement standards analyzed to calculate a single MU; b) select data from one specified measurement standard (e.g., a concentration at or near a legal specification); or c) calculate the quantity value for Type A data at each measurement standard concentration.
Calculation of the Quantity Value for Type A Data | 5.4.2.3.2.2.2 | Requirement | If a consistent variance is not demonstrated, breath alcohol programs shall: a) utilize the Type A data from the measurement standard producing the largest variance; b) perform an in-depth evaluation to determine where the variance changes occur across the calibration range and establish an appropriate uncertainty to report based on where these variance changes occur; or c) calculate the MU at each measurement standard concentration.
Calculation of the Quantity Value for Type A Data | 5.4.2.3.2.3 | Requirement | Multiple instruments. The calibration method and instrument make/model shall be the same to calculate MU by combining data from multiple breath alcohol measuring instruments.
Minimum Requirements for Type B Evaluations | 5.4.3 | Section Title
Minimum Requirements for Type B Evaluations | 5.4.3.1 | Informational Text | Components requiring a Type B evaluation may include uncertainty associated with a certified reference material, uncertainty of a reference material, and/or uncertainty from equipment calibration (e.g., balance, volumetric flask, pipette, barometer, or thermometer).
Minimum Requirements for Type B Evaluations | 5.4.3.2 | Requirement | When considering which components to include in the Type B evaluations, testing laboratories and breath alcohol programs shall: a) consider all components that are not accounted for in a Type A evaluation; b) ensure components are evaluated according to the assumed distribution of the quantity value; and c) account for all identified and significant systematic bias (see 5. 6. 2).
Establishing a quantity value for Type B evaluations | 5.4.4 | Section Title
Establishing a quantity value for Type B evaluations | 5.4.4.1 | Requirement | For component(s) used in the preparation of a calibrator or measurement standard, the components shall be quantified individually or as a group.
Establishing a quantity value for Type B evaluations | 5.4.4.1.1 | Requirement | If evaluating uncertainty over the full calibration range, testing laboratories and breath alcohol programs shall use the largest standard deviation calculated.
Establishing a quantity value for Type B evaluations | 5.4.4.1.2 | Requirement | If evaluating the uncertainty for multiple concentration ranges, testing laboratories and breath alcohol programs shall use the largest standard deviation calculated for each concentration range, respectively.
Establishing a quantity value for Type B evaluations | 5.4.4.1.3 | Requirement | If evaluating the uncertainty at each calibrator or measurement standard concentration separately, testing laboratories and breath alcohol programs shall use the value for the applicable calibrator or measurement standard. NOTE 1 If the test or calibration method includes the preparation of multiple calibrators or measurement standards, the individual components may be quantified individually across all calibrator concentrations (e. g., a single component quantity value can be used for the pipette uncertainty that adequately covers the pipettes used to prepare all calibrator concentrations) and then 5. 4. 4. 1. 1 or 5. 4. 4. 1. 2 above may be applied. Alternatively, the components may be quantified as a group for each calibrator concentration and then 5. 4. 4. 1. 1 through 5. 4. 4. 1. 3 applied. NOTE 2 Depending on the measurement process, these components related to calibrator preparation, typically requiring a Type B evaluation, may be accounted for by ongoing control data (Type A).
Step 4—Convert Quantities to Standard Uncertainties | 5.5 | Section Title
General | 5.5.1 | Requirement | The testing laboratory or breath alcohol program shall quantify all uncertainty components as a standard uncertainty of the quantity values and in the same measurement unit or in a measurement unit relative to the quantity values.
Type A Evaluations | 5.5.2 | Section Title
Type A Evaluations | 5.5.2.1 | Requirement | If not already presented as a standard uncertainty, the quantity shall be divided by the appropriate coverage factor (k) to convert to a standard uncertainty.
Type B Evaluations | 5.5.3 | Section Title
Type B Evaluations | 5.5.3.1 | Requirement | If not reported by the manufacturer as a standard uncertainty, the testing laboratory or breath alcohol program shall use the appropriate probability density function for the component to compute one standard deviation or relative standard deviation associated with the specified distribution.
Type B Evaluations | 5.5.3.2 | Requirement | If reported by the manufacturer as an expanded uncertainty (U), the testing laboratory or breath alcohol program shall divide by the appropriate coverage factor (k) to arrive at a standard uncertainty.
Step 5—Calculate the Combined Standard Uncertainty | 5.6 | Section Title
General | 5.6.1 | Requirement | The testing laboratory or breath alcohol program shall calculate the combined standard uncertainty using the uncertainty contributors’ quantity values, utilizing the root sum of the squares formula or the Monte Carlo method.
General | 5.6.1.1 | Requirement | A justification shall be documented if any uncertainty component is excluded from the combined standard uncertainty.
Evaluation of Bias | 5.6.2 | Section Title
Evaluation of Bias | 5.6.2.1 | Informational Text | Measurement accuracy encompasses both precision and bias. A measurement is more accurate when it has less bias and greater precision. The GUM states, “It is assumed that the result of a measurement has been corrected for all recognized significant systematic effects and that every effort has been made to identify such effects. ” An evaluation of bias may not always be possible as one or more controls prepared with metrological traceability, having a known reference value and uncertainty, is required to evaluate bias.
Evaluation of Bias | 5.6.2.2 | Requirement | Bias evaluation shall be performed whenever possible.
Evaluation of Bias | 5.6.2.3 | Requirement | The general approach to bias evaluation shall: a) Determine if bias is present by comparing measurement standard or control data to reference values with established metrological traceability; b) Calculate the combined uncertainty (u) without considering the relevant bias; and c) Compare the bias with the combined standard uncertainty. 1) Where the bias is less than the combined standard uncertainty, bias<uc, the bias is viewed as insignificant and may be neglected or included as a component in the uncertainty evaluation. 2) Where the bias is greater than or equal to the combined standard uncertainty, bias≥uc, it is considered significant and additional action shall be taken; see 5.6.2.4 and 5.6.2.5.
Evaluation of Bias | 5.6.2.4 | Requirement | Testing laboratories shall address significant bias in one of the following ways: a) modify the method to reduce the bias until it is no longer significant and the expanded uncertainty of the method remains fit for purpose; b) correct the measurement result for the bias, including the uncertainty of the correction in the evaluation of uncertainty. Both the observed measurement result and the corrected measurement result with the MU shall be reported; c) report the measurement result and the expanded MU with bias included; or d) report the observed measurement result, the MU, and the bias.
Evaluation of Bias | 5.6.2.5 | Requirement | Breath alcohol programs shall address significant bias in one of the following ways: a) modify the method to reduce the bias until it is no longer significant and the expanded uncertainty of the method remains fit for purpose; b) report the measurement result and the expanded MU with bias included; or c) report the observed measurement result, the MU, and the bias.
Step 6—Calculate the Expanded Uncertainty | 5.7 | Section Title
Step 6—Calculate the Expanded Uncertainty | 5.7.1 | Requirement | A coverage factor (k) shall be determined using a Student’s t-distribution based on the degrees of freedom (n-1) to provide the desired confidence level.
Step 6—Calculate the Expanded Uncertainty | 5.7.2 | Requirement | The minimum coverage probability for all quantitative test results and calibration results shall be 95.45 %.
Step 7—Evaluate the Expanded Uncertainty | 5.8 | Section Title
Step 7—Evaluate the Expanded Uncertainty | 5.8.1 | Requirement | A determination of whether the calculated measurement uncertainty is acceptable shall be made by the testing laboratory or breath alcohol program.
Step 7—Evaluate the Expanded Uncertainty | 5.8.2 | Requirement | The evaluation of acceptance, as applicable, shall consider: a) stakeholder interests; b) legal requirements; c) the relationship between the reported test or calibration quantitative value and the expanded MU; particular consideration shall be taken around the LLOQ/LOD; and EXAMPLE An expanded MU of 0.01 ng/mL for a method with an LLOQ of 0.01 ng/mL would prompt the testing laboratory or breath alcohol program to reevaluate the LLOQ. d) the relationship between the control limits for the method and the expanded measurement uncertainty. EXAMPLE Control limits of ± 20 % for a method with expanded MU of 10 % (95.45 % coverage probability). For any single analytical batch, this control limit would allow a variation of up to 20 %, which exceeds the stated expanded MU for the method and would prompt the testing laboratory or breath alcohol program to reevaluate the control limits.
Step 8—Report the Expanded Uncertainty | 5.9 | Section Title
Step 8—Report the Expanded Uncertainty | 5.9.1 | Requirement | For testing laboratories, MU reporting shall be in accordance with ANSI/ASB Standard 053, Standard for Reporting in Forensic Toxicology.
Step 8—Report the Expanded Uncertainty | 5.9.2 | Requirement | For breath alcohol programs, the MU shall be reported as part of the calibration result.
Step 8—Report the Expanded Uncertainty | 5.9.3 | Informational Text | When the MU is reported:
Step 8—Report the Expanded Uncertainty | 5.9.3.1 | Requirement | For testing laboratories, the MU shall be reported as an expanded uncertainty and include the coverage probability.
Step 8—Report the Expanded Uncertainty | 5.9.3.2 | Requirement | For breath alcohol programs, the MU shall be reported as an expanded uncertainty and include the coverage factor, k, and the coverage probability.
Step 8—Report the Expanded Uncertainty | 5.9.3.3 | Requirement | The measurement result shall include the measured quantity value, y, along with the associated expanded uncertainty, U. It should be reported as y ± U, where U is consistent with the units of y. Specific applications may warrant using a different format than y ± U.
Step 8—Report the Expanded Uncertainty | 5.9.3.4 | Recommendation | The expanded uncertainty should be reported to at most 2 significant figures unless the testing laboratory or breath alcohol program has a documented rationale to report beyond 2 significant figures.
Step 8—Report the Expanded Uncertainty | 5.9.3.5 | Requirement | Rules for rounding the expanded uncertainty shall be defined by the testing laboratory or breath alcohol program.
Step 8—Report the Expanded Uncertainty | 5.9.3.6 | Requirement | The rounded expanded uncertainty shall be reported using the same number of decimal places as the measurement result unless a legal specification specifies how the result will be reported. Rules for rounding or truncating the measurement result shall be defined by the testing laboratory or breath alcohol program.
Step 8—Report the Expanded Uncertainty | 5.9.3.7 | Requirement | Testing laboratories shall report the respective measurement uncertainty for each analyte within a method. NOTE Combining the MU across multiple analytes or methods would lead to an overestimation of the MU, which does not meet the intent of a measurement uncertainty evaluation.
Step 8—Report the Expanded Uncertainty | 5.9.3.8 | Requirement | For testing laboratories, if a significant bias is identified and the action taken is as described in 5.6.2.3 b) or c), this shall be clearly communicated.
Periodic Evaluation of Measurement Uncertainty. | 6 | Section Title
Periodic Evaluation of Measurement Uncertainty. | 6.1 | Requirement | The testing laboratory or breath alcohol program shall set the interval for reviewing and recalculating a method’s MU and shall retain records supporting the decision.
Periodic Evaluation of Measurement Uncertainty. | 6.2 | Requirement | For both Type A and Type B uncertainty components included in the MU calculation, the decision shall consider: a) the frequency with which one of the components changes; b) the frequency with which the testing or calibration method is performed; c) the magnitude of a change in a component in relationship to the calculated MU; d) subsequent sources of Type A data (e.g., changes to personnel, additional instrumentation); e) a change in the measurement process; and f) any testing laboratory or breath alcohol program administrative decision such as a set time interval.
Periodic Evaluation of Measurement Uncertainty. | 6.3 | Requirement | Any recalculation of the measurement uncertainty shall meet all requirements of this standard.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed | Clause involves options; this option was not chosen | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Forensic Document Examination
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics