## Instructions for Use
AAFS Standards Resources and Training
This checklist is made possible through the following financial assistance award 70NANB24H078 awarded to AAFS from U.S. Department of Commerce, National Institute of Standards and Technology Each checklist is specific to a version of a standard, which is hyperlinked in the checklist. Not all sections of a standard are included in a checklist. A checklist is a supplementary tool and should not be used independently. To properly use a checklist, reference the complete standard. Go to https://www.nist.gov/organization-scientific-area-committees-forensic-science/access-standards for information on how to access published standards. A checklist does not provide interpretation guidance. If you have questions related to interpretation, please contact the standard developing organization (SDO). The SDO may have standard specific training available.
A checklist can be used for one or both of the following purposes: • to evaluate the implementation of a standard, or • to audit (internal or external) conformance with requirements in the standard and an FSSP’s management system
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions and an on-demand training webinar.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E3406-25e1
ANSI/ASTM E3406-25e1
Standard Guide for Microspectrophotometry in Forensic Fiber Analysis
Discipline: | Trace Evidence
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Sample Preparation | 6 | Section Title
Sample Preparation | 6.1 | Informational Text | The general collection, handling, and tracking of samples are to meet or exceed the requirements of Guide E1459 and Practice E1492.
Sample Preparation | 6.2 | Informational Text | The work area and tools used for the preparation of samples are cleaned prior to use.
Sample Preparation | 6.3 | Informational Text | Known and questioned samples are mounted and prepared in the same manner. Individual fibers from a sample are dispersed on the slide in order to minimize interference and the effects of photobleaching.
Sample Preparation | 6.4 | Section Title | Transmittance Measurements:
Sample Preparation | 6.4.1 | Informational Text | The fiber(s) is mounted on a microscope slide under a coverslip in an appropriate medium.
Sample Preparation | 6.4.1.1 | Informational Text | A microscope slide with transmittance characteristics appropriate for the region of the spectrum being analyzed is used. Glass slides and coverslips are suitable for measurements in the visible and NIR portions of the spectrum. Typical soda-lime glass slides absorb light in the UV region, therefore samples to be analyzed in the UV region of the spectrum are mounted between quartz or fused silica slides and coverslips. If performing fluorescence measurements, it is important that the slide and coverslip have low or no inherent fluorescence.
Sample Preparation | 6.4.1.2 | Informational Text | A mounting medium that is compatible with the sample (that is, will not dissolve the fiber or dye) and the spectral range being investigated is selected. Mounting media include, but are not limited to, water, xylene, xylene substitutes, glycerol, and refractive index oils (n = 1.52 or 1.66 are common). When performing fluorescence measurements, a mounting medium with low or no inherent fluorescence is selected. For example, glycerol is a suitable mounting medium when analyzing a sample in the UV, visible, and NIR regions and when performing fluorescence measurements.
Performance Checks | 7 | Section Title
Performance Checks | 7.1 | Informational Text | Prior to use of the instrument, the microscope, illumination sources, and spectrometer are turned on and allowed to stabilize. This is performed in accordance with the instrument manufacturers’ instructions or laboratory experience, whichever yields consistent results.
Performance Checks | 7.2 | Informational Text | Checking instrument performance verifies that an instrument is operating within required standards. It is essential to demonstrate wavelength and absorbance/photometric accuracy through a performance check, such as that described in Practice E275.
Performance Checks | 7.2.1 | Informational Text | A performance check is conducted each day of use, prior to analysis.
Performance Checks | 7.2.2 | Informational Text | A performance check is conducted after any maintenance or power outages, prior to analysis.
Performance Checks | 7.2.3 | Informational Text | A similar configuration is used each time a performance check is conducted on the system.
Performance Checks | 7.3 | Informational Text | Records of all performance checks are maintained. A historical record of this data provides a mechanism for evaluating system performance and allows an operator to monitor system trends and deterioration.
Performance Checks | 7.4 | Section Title | Performance Check Parameters:
Performance Checks | 7.4.1 | Recommendation | Wavelength Accuracy—Wavelength accuracy over the measured range is checked with the aid of manufacturer-recommended filters (for example, holmium-, erbium-, or didymium oxide) The resolution used during the wavelength accuracy checks should be the same as that used in casework and consistent for each wavelength accuracy check. Transmittance is used for these measurements.
Performance Checks | 7.4.2 | Informational Text | Photometric Accuracy—The photometric response of the system is checked to ensure linearity using manufacturer-recommended neutral density filters. A typical set of neutral density calibration filters could include some or all of the following filters: 0.1, 0.5, 1.0, 2.0, 2.5, and 3.0 absorbance units.
Performance Checks | 7.4.3 | Informational Text | Filters are calibrated according to the manufacturer’s recommendations.
Performance Checks | 7.5 | Informational Text | Fluorescence Emission—Fluorescence emissions are checked with materials known to fluoresce (for example, optically-brightened cotton, fluorescent fingerprint powder, SRM 2940) according to the manufacturer’s recommendations.
Instrument and Scanning Parameters | 8 | Section Title
Instrument and Scanning Parameters | 8.1 | Informational Text | MSP instruments can vary and specific details on the operation and system parameters can be found in the manufacturer’s manuals and guides.
Instrument and Scanning Parameters | 8.2 | Section Title | Microscope Parameters:
Instrument and Scanning Parameters | 8.2.1 | Requirement | Illuminator—An illumination source appropriate for the analysis being conducted is selected. The illuminator shall have sufficient intensity across the entire wavelength range of interest so as to provide a spectrum with an acceptable signal-to-noise ratio according to the manufacturer’s recommendation. Tungsten, halogen, and xenon lamps are commonly used for visible and NIR analysis. Xenon lamps are frequently used for UV analysis and mercury lamps are used for fluorescence excitation. While LED illuminators are available over much of the spectrum, they are of little utility for microspectrophotometry due to their lower intensity and limited spectral range.
Instrument and Scanning Parameters | 8.2.1.1 | Informational Text | Background, system, and reference transmittance spectra can be used to monitor illuminator performance and warn of unsuitable system alignment.
Instrument and Scanning Parameters | 8.2.1.2 | Informational Text | Illumination Centration—Slight adjustments to the position of the bulb can serve to increase or decrease the emission over specific regions of the spectrum. For example, it is possible to maximize UV illumination, often at the expense of some light in the visible wavelengths. Generally, the slight loss of intensity in the visible region is not problematic due to the high intensity of modern bulbs.
Instrument and Scanning Parameters | 8.2.1.3 | Recommendation | Illumination Intensity—For some illuminators, this can be a fixed parameter. When the voltage of an illuminator is adjustable, it should be held fixed following the photometric intensity performance check (see 7.4.2).
Instrument and Scanning Parameters | 8.2.2 | Informational Text | Field Diaphragm—With the specimen in focus, the edges of the field diaphragm are brought into view and then sharply focused by adjusting the substage condenser. The field diaphragm is then opened so that its edges are either just outside the collection aperture or just beyond the field of view to minimize stray light. The focus and size of the field diaphragm is readjusted when the objective is changed.
Instrument and Scanning Parameters | 8.2.3 | Informational Text | Substage Aperture (that is, Condenser Iris)—The substage aperture is opened until the desired image contrast is obtained. As adjustment of this aperture impacts the amount of light reaching the detector, this aperture setting can be different from that used to produce an ideal image. In some instances, it is desirable to further increase the opening of the substage aperture to allow more light to reach the detector. The aperture is adjusted before collecting the background spectrum and kept in a fixed position between background and sample spectra collection and between samples when they are being compared. The appropriate aperture level is typically that which produces an emission intensity for the most intense emission peak of no more than approximately 80 % of the detector saturation value. It is critical that the detector is not saturated anywhere over the region being measured.
Instrument and Scanning Parameters | 8.2.4 | Informational Text | Objective—An objective that permits visualization of the fiber to be analyzed is selected. A balance between the objective magnification and size of the measuring aperture is selected by the FSP. Typically, measurements are made using objectives between 10× and 50×. Quartz optics are required for measurements made in the UV region. Once the appropriate objective is selected, all samples being compared are measured at a fixed magnification.
Instrument and Scanning Parameters | 8.2.5 | Recommendation | Measuring Aperture—If the MSP system is equipped with variable or multiple collection apertures, the same aperture is used for the measurement of all samples being compared. In general, the largest aperture that will remain within the boundaries of the sample area to be measured is selected. Analysis of the edge of a sample should be avoided due to edge effects that could impact the spectrum. An oversized aperture (one that extends beyond the boundary of the sample) is undesirable in transmittance measurements as it decreases the photometric accuracy of the measurement. In fluorescence emission, the spectrum is not the result of a ratio to a reference scan; the strength of the signal is determined by absolute counts. Therefore, an oversized aperture when analyzing fibers for fluorescence emission can be used to increase the signal reaching the detector. As the background is black, there is no significant increase in noise to detract from the quality of the collected data.
Instrument and Scanning Parameters | 8.2.6 | Informational Text | Filter Cube—Some systems have a filter cube turret. Filter cubes can include a blank cube for transmission measurements, a mirror for reflection, and combinations of excitation and barrier/emission filters for fluorescence measurements. A filter cube appropriate to the measurement being made is chosen. All samples being compared are measured under a given configuration with the same filter cube in place.
Instrument and Scanning Parameters | 8.2.7 | Informational Text | Phototube Diverter—Some systems have an adjustable phototube that permits light to be diverted to the ocular, spectrometer/camera, or split between the ocular and spectrometer/camera. Spectral artifacts (for example, interference fringes on the baseline) could be visible if the diverter is in the split position. The diverter is set so that all light is directed to the spectrometer during measurements to ensure the highest possible signal-to-noise ratio.
Instrument and Scanning Parameters | 8.3 | Section Title | Spectrometer Parameters:
Instrument and Scanning Parameters | 8.3.1 | Informational Text | Detector—While the detector is not an adjustable variable post-purchase, it plays a critical role during data collection. Two main detector types are available for use with MSP systems.
Instrument and Scanning Parameters | 8.3.1.1 | Informational Text | Photomultiplier tube (PMT) detectors consist of a photocathode, held at a positive potential, and a series of dynodes with successively lower potential which amplify the signal and convert photons of light into electrical energy. PMTs are sensitive, provide a high signal-to-noise ratio, and have good spectral resolution. PMTs are typically single channel detectors and are generally used in scanning spectrophotometers in conjunction with a grating. The sample is scanned by stepping through wavelengths to create a spectrum point by point.
Instrument and Scanning Parameters | 8.3.1.2 | Informational Text | Semiconductor detectors are composed of a monochromator fitted with a diffraction grating and an array detector (for example, charge-coupled devices (CCD)) that acts as the photosensitive device. A CCD detector generally has a lower signal-to-noise ratio when compared to a PMT detector, but measurement time is drastically reduced because of simultaneous detection of the full spectral range. For this reason, CCD detectors are far more common than PMT detectors. The CCD detector’s resolution will depend on the number of pixels in the array, the dispersion and line spacing of the grating, and the distance between the grating and the array.
Instrument and Scanning Parameters | 8.3.2 | Informational Text | Wavelength Range—A wavelength range is selected that is appropriate to the desired range of measurement, instrument capabilities, sample, and sample preparation conditions as discussed in Section 6. This range typically falls between 190 nm and 1100 nm (transmittance UV-Vis-NIR).
Instrument and Scanning Parameters | 8.3.3 | Informational Text | Resolution—The resolution is predominantly defined by the grating and slit size. For most commercial MSPs, the grating and slit size are fixed, which results in a fixed maximum resolution. For CCD detectors, pixel size and spacing also affect resolution.
Instrument and Scanning Parameters | 8.3.4 | Informational Text | Integration Time—Most MSP systems allow the user to define the integration time. Some software packages have a built-in functionality that automatically adjusts the integration time. When this is not available, the integration time is set such that the detector electronics are not saturated.
Instrument and Scanning Parameters | 8.3.4.1 | Informational Text | To manually optimize the integration time, a background spectrum is collected and the highest peak is checked to ensure that it is not saturated. If the detector is saturated (the light intensity is too high), the integration time is reduced. Conversely, if the light intensity is too low, the integration time is increased. In either instance, recollect the dark and background spectra before your sample.
Instrument and Scanning Parameters | 8.3.5 | Requirement | Number of Scans—Most MSP systems allow the user to define the number of scans to average for a single spectrum. Select the number of scans that yields the acceptable signal-to-noise ratio according to the manufacturer’s recommendation. Generally, the signal-to-noise ratio is proportional to the square root of the number of scans (for example, if the number of scans is quadrupled, the signal-to-noise ratio is increased by a factor of 2).
Instrument and Scanning Parameters | 8.4 | Section Title | Determine Polarization in the System:
Instrument and Scanning Parameters | 8.4.1 | Recommendation | The set-up of some MSP microscopes results in polarization of the transmitted light, therefore each system should be assessed to determine the effects of possible polarization. A moderately to highly dichroic fiber is analyzed at different orientations (for example, 0 degrees, 45 degrees, and 90 degrees) and the resultant spectra compared. If the orientation is found to affect spectral results, casework samples should be oriented in the same direction during analyses to minimize the effects of the possible polarization or spectra should be collected from a wide range of fiber orientations.
Sample Analysis | 9 | Section Title
Sample Analysis | 9.1 | Informational Text | Prior to sample analysis, performance checks are conducted in accordance with Section 7.
Sample Analysis | 9.2 | Recommendation | The illumination field, measuring aperture, and magnification for the sample under investigation are optimized. Experimental conditions and instrument settings (for example, objective, aperture size, illumination intensity, number of scans, resolution) should be identical for compared samples.
Sample Analysis | 9.2.1 | Informational Text | In order to reduce sample degradation (that is, photodegradation), the amount of time the sample is illuminated when it is not being actively analyzed is minimized (9, 13, 14). The potential for sample degradation can be assessed on known samples prior to analyzing questioned fibers. Photodegradation during sample analysis is typically due to the UV wavelengths from the MSP light source.
Sample Analysis | 9.2.1.1 | Informational Text | A UV blocking filter can be placed between the condenser and the light source while spectra are not being collected.
Sample Analysis | 9.2.2 | Informational Text | Consistent sample conditions are especially critical when collecting fluorescence data because the amplitude of the resulting spectrum is directly correlated with the collection conditions (that is, there is no reference scan to ratio against).
Sample Analysis | 9.2.3 | Informational Text | Consistent or similar sample orientations are preferred when comparing samples. This can be accomplished most easily via a rotating stage, but can also be done by ensuring fibers are mounted in the same general orientation.
Sample Analysis | 9.2.3.1 | Informational Text | Dichroic fibers in different orientations can yield spectra with significant differences when using polarized light. Collecting spectra from a dichroic fiber in both the parallel and perpendicular orientations provides more data for comparison.
Sample Analysis | 9.3 | Section Title | Collection of Reference Spectra
Sample Analysis | 9.3.1 | Section Title | Dark Scans:
Sample Analysis | 9.3.1.1 | Informational Text | A dark scan is a reference spectrum collected when the light from the microscope is blocked from the detector and is a measurement of instrument noise.
Sample Analysis | 9.3.1.2 | Informational Text | A dark scan is collected prior to the analysis of each new microscope slide or sample preparation.
Sample Analysis | 9.3.1.3 | Informational Text | A dark scan is required for transmittance, reflectance, and fluorescence emission spectra.
Sample Analysis | 9.3.2 | Section Title | Background Scans:
Sample Analysis | 9.3.2.1 | Informational Text | A background scan is a spectrum that measures the light transmitting/absorbing effect of all the system components (that is, light source, optics, microscope slide, cover slip, and mounting medium) without the sample of interest in the light path. The background scan is also called the “reference scan.”
Sample Analysis | 9.3.2.2 | Informational Text | A background scan is required for transmittance and reflectance measurements, but is not relevant to fluorescence emission spectra.
Sample Analysis | 9.3.2.3 | Informational Text | For transmittance measurements, a new background scan is collected for each new microscope slide and coverslip preparation, aperture size, or instrument configuration. In general, a background scan collected before every sample scan will best compensate for the effects of adjusting the fine focus on the microscope and of heterogeneities in the mounting medium.
Sample Analysis | 9.4 | Section Title | Collection of Sample Spectra:
Sample Analysis | 9.4.1 | Informational Text | For fibers with differences in thickness due to cross-sectional shape (for example, trilobal, triangular, flattened), sampling areas are chosen in the compared fibers with similar thicknesses and orientations (for example, through a single lobe of a trilobal fiber or the flat area of a flattened/ribbon fiber).
Sample Analysis | 9.4.2 | Recommendation | Multiple spectra should be collected from each uniformly-colored sample, each representing different sample areas. For fibers that are not uniformly-colored (for example, cotton and other natural fibers), a larger number of spectra is recommended. The number of replicate analyses can be adjusted in an effort to capture the variation present within the sample. Small sample size or poor sample conditions could limit the acquisition of multiple spectra.
Sample Analysis | 9.4.3 | Informational Text | Known sample(s) representative of the variation in color within the textile are selected. Differences could arise in measurements of fiber samples from the same garment or textile because of differences in inherent variation in color between fibers, weathering (for example, sunlight exposure), spot staining/bleaching, or repaired areas (for example, use of a fabric marker to cover a discolored area, application of a patch).
Spectral Comparison and Interpretation | 10 | Section Title
Spectral Comparison and Interpretation | 10.1 | Informational Text | Spectral comparisons are to be conducted between spectra collected using the same sample preparation methods, similar sample characteristics (for example, color intensity, thickness, orientation), and the same instrumental parameters.
Spectral Comparison and Interpretation | 10.2 | Informational Text | The spectra are compared based on the observation of any spectral differences, or lack thereof, between the sets of microspectrophotometric data.
Spectral Comparison and Interpretation | 10.2.1 | Informational Text | The sample comparison begins with the examination of the whole spectrum, followed by critical examination of each specific peak. The comparison includes examination of peak shape, minima, maxima, inflection points, troughs, shoulders, relative peak intensities, and the curves or slopes between peaks.
Spectral Comparison and Interpretation | 10.3 | Informational Text | Spectral overlay is a method for comparing data where the presence or absence of peaks, peak shapes, and relative intensities are all considered in the evaluation as to whether exclusionary differences exist between compared samples.
Spectral Comparison and Interpretation | 10.3.1 | Informational Text | Spectral comparisons can be conducted with the spectra displayed in percent reflectance, percent transmittance, or absorbance formats. Certain information, however, is observed more readily in one format or the other. Absorbance is better for observing differences in darkly colored fibers. Transmittance is better for observing differences in lightly colored fibers. At a minimum, spectra are presented on the same x-axis scale when providing overlays or performing comparisons.
Spectral Comparison and Interpretation | 10.3.1.1 | Recommendation | Mean value spectra (that is, averaged) can be generated from replicate scans of each sample. Spectra are typically averaged and then compared to the averaged spectra of another item. Mean spectra should be calculated from absorbance data.
Spectral Comparison and Interpretation | 10.3.1.2 | Recommendation | Plots of standard deviation spectra (calculated from multiple spectra collected from a given sample) can also provide a useful point of comparison. Standard deviation curves can be useful for estimating the known sample variation range; however, comparisons based upon standard deviation spectra use intensity as a criterion for comparison. Standard deviations should be calculated from absorbance spectra.
Spectral Comparison and Interpretation | 10.3.1.3 | Informational Text | First and second derivative functions of the spectra can also assist in identifying inflection points and aid in the discrimination of samples. Effective use of derivative functions requires that spectra have high signal-to-noise ratios. Spectral derivative calculations are conducted on absorbance data.
Spectral Comparison and Interpretation | 10.4 | Informational Text | Sample limitations (for example, small samples, dirty samples, color intensity variations) and instrumental limitations (for example, sampling size) are considered when assessing differences between spectra.
Spectral Comparison and Interpretation | 10.4.1 | Informational Text | Possible reasons for spectral differences include dissimilar sample characteristics, heterogeneity, contribution from extraneous materials, or origination from different source materials. Collecting spectra from additional samples can provide supplemental data to assist in assessing such differences. Slight differences in peak heights can indicate differences in dye concentrations, light exposure (fading), or dye uptake.
Spectral Comparison and Interpretation | 10.4.2 | Informational Text | Some spectral differences are subtle and visually difficult to discern. In these instances, chemometric analysis could help assess compared samples. It is noted that consensus has not been reached in the relevant scientific literature on the most appropriate chemometrics for application to microspectrophotometric data.
Spectral Comparison and Interpretation | 10.5 | Informational Text | If suitable spectra are produced, comparisons can provide information regarding the potential relationship between the sources of the samples.
Spectral Comparison and Interpretation | 10.5.1 | Informational Text | Distinguishable Sources—When exclusionary differences are observed between compared spectral features, the sources of the samples are considered distinguished by microspectrophotometry. Exclusionary differences in microspectrophotometric spectral comparisons: (1) are outside the variability of spectra originating from the same source; and (2) cannot be explained by considerations such as sample heterogeneity, contamination, different sample conditions, or different sample histories.
Spectral Comparison and Interpretation | 10.5.2 | Informational Text | Indistinguishable Sources—When no exclusionary differences are observed between compared spectral features, the sources of the samples are considered indistinguishable by microspectrophotometry. Differences that are not considered exclusionary: (1) are within the variability of spectra originating from the same source; or (2) can be explained by considerations such as sample heterogeneity, contamination, different sample conditions, or different sample histories. If no exclusionary differences are observed in a microspectrophotometric spectral comparison, samples can be analyzed by other analytical techniques to provide additional information about the potential relationship between the sources of the samples.
Spectral Comparison and Interpretation | 10.6 | Informational Text | Microspectrophotometric spectral comparison is one part of a multi-analytical comparative approach. Microspectrophotometric data alone can be used to distinguish the sources of compared samples, but they are not used independent of data obtained from other analytical techniques to reach an overall opinion regarding the potential relationship between the sources of the samples. An overall opinion that sources are indistinguishable is only reported when no exclusionary differences are observed in any of the analytical techniques that were applied.
Examination Documentation | 11 | Section Title
Examination Documentation | 11.1 | Informational Text | The documentation discussed below is only in regard to the instrumental analysis of fibers by MSP and the interpretation of the resulting spectra. FSP notes and included data are to be of sufficient detail to allow an independent FSP to understand and evaluate all work performed, analyze and interpret the data, and reach independent opinions.
Examination Documentation | 11.1.1 | Informational Text | Instrumental data used for interpretations is included in the case notes.
Examination Documentation | 11.1.2 | Informational Text | The details necessary to support and understand the interpretations made from each spectral comparison are recorded (refer to Practice E620).
Examination Documentation | 11.2 | Informational Text | A description of the evidence analyzed by MSP, the method of sample preparation (including any mounting medium used), the analytical instrumentation used, mode of operation (transmission, fluorescence, etc.), and its optimized operating parameters (for example, aperture size, objective, scan/spectrum averaging, spectral resolution, fluorescence filter cube) is included in the case notes, case record, or otherwise recorded in accordance with laboratory procedures.
Examination Documentation | 11.3 | Informational Text | Spectra are provided either in color or in a format such that spectra from various samples plotted together can be attributed to a legend when viewed in grayscale.
Examination Documentation | 11.4 | Informational Text | If chemometric methods are applied, the data analysis method(s) and all information (for example, software name and version, confidence intervals) necessary to review the result are recorded.
Examination Documentation | 11.5 | Informational Text | Refer to Practices E1492 and E620, and ISO/IEC 17025 for further guidance.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed | Clause involves options; this option was not chosen | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Forensic Document Examination
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics