## Instructions for Use
AAFS Standards Resources and Training
This checklist is made possible through the following financial assistance award 70NANB24H078 awarded to AAFS from U.S. Department of Commerce, National Institute of Standards and Technology Each checklist is specific to a version of a standard, which is hyperlinked in the checklist. Not all sections of a standard are included in a checklist. A checklist is a supplementary tool and should not be used independently. To properly use a checklist, reference the complete standard. Go to https://www.nist.gov/organization-scientific-area-committees-forensic-science/access-standards for information on how to access published standards. A checklist does not provide interpretation guidance. If you have questions related to interpretation, please contact the standard developing organization (SDO). The SDO may have standard specific training available.
A checklist can be used for one or both of the following purposes: • to evaluate the implementation of a standard, or • to audit (internal or external) conformance with requirements in the standard and an FSSP’s management system
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions and an on-demand training webinar.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E3501-25
ANSI/ASTM E3501-25
Standard Guide for Forensic Analysis of Geological Materials by Scanning Electron Microscopy and Energy Dispersive X-Ray Spectrometry
Discipline: | Trace Evidence
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Specimen Preparation Techniques | 5 | Section Title
Specimen Preparation Techniques | 5.1 | Informational Text | The specimen preparation procedures selected depend upon the goals of the forensic examination, the type of material (consolidated versus unconsolidated, particle sizes, amount of material available), and any other methods of analysis to be applied before or after SEM/EDS.
Specimen Preparation Techniques | 5.2 | Section Title | General Requirements and Recommendations:
Specimen Preparation Techniques | 5.2.1 | Requirement | Use specimen sample preparation practices that minimize the possibility of cross-contamination. Prepare questioned and known items on separate SEM specimen mounts (for example, stubs) (refer to Guide E3457). NOTE 1—The nature of the samples affects precautionary measures recommended to prevent cross contamination. Approaches for minimizing cross contamination include: preparing questioned and known samples in separate locations, storing prepared specimens within covered containers to minimize dust accumulation, and placing questioned and known items within the SEM/EDS instrument at separate times. Specimens of limited quantity and small particle size not embedded in mounting material require the greatest measures to mitigate potential cross contamination (refer to Guide E3457).
Specimen Preparation Techniques | 5.2.2 | Informational Text | Exposure of a sticky mount during sample preparation can provide a control specimen allowing tests for possible airborne contamination.
Specimen Preparation Techniques | 5.2.3 | Requirement | Document the location of the materials to be analyzed on the SEM mount when more than one specimen or sub-sample is placed on the same mount. Documentation methods can include sketches, photographs, captured video images, or index (fiducial) marks on the mount. Label specimens with clear and unique identifiers.
Specimen Preparation Techniques | 5.2.4 | Informational Text | Specimen preparation is commonly carried out with the aid of a stereomicroscope.
Specimen Preparation Techniques | 5.2.5 | Requirement | Once the SEM mounts are prepared, protect them from surface abrasion and atmospheric dust deposition with a protective lid.
Specimen Preparation Techniques | 5.2.6 | Informational Text | Storing the prepared specimens in a vacuum chamber or desiccator reduces SEM chamber evacuation time.
Specimen Preparation Techniques | 5.2.7 | Informational Text | Charging can be mitigated by applying a surface coating (for example, carbon or metal) or adhering to the mount with a conductive paint (for example, silver or carbon paint). Alternatively uncoated specimens can be analyzed using variable pressure mode (see 3.1.8) or using low accelerating voltage or beam current. NOTE 2—Variable pressure mode increases the width of the spot size and can reduce the EDS signal. The beam scatter could result in contributions to the EDS spectrum from the mounting material or other material that is in proximity. The wider spot size could also reduce image quality and degrade spatial resolution. The benefits of retaining uncoated samples are balanced against the reduced data quality and reduced spatial resolution. NOTE 3—Lower beam current can reduce detection limits and lower accelerating voltage can affect the excitation of elements.
Specimen Preparation Techniques | 5.2.7.1 | Informational Text | Specimen coating is commonly accomplished using a vacuum evaporator or sputter coater. NOTE 4—Metal (for example, gold, palladium, platinum) coatings can interfere with X-ray peak assignments (for example, determination of phosphorus and zirconium). If coating is required, carbon is preferred to metals when EDS analysis is performed.
Specimen Preparation Techniques | 5.2.8 | Requirement | If the intent of the analysis is the comparison of two or more samples, use the same specimen preparation technique for all samples to be compared.
Specimen Preparation Techniques | 5.3 | Section Title | Unconsolidated Materials—Soils and Sediments:
Specimen Preparation Techniques | 5.3.1 | Section Title | Separation of Unconsolidated Materials:
Specimen Preparation Techniques | 5.3.1.1 | Informational Text | Particle separation can be beneficial prior to SEM/EDS analysis.
Specimen Preparation Techniques | 5.3.1.2 | Informational Text | Processing of unconsolidated materials can include grain washing, particle size fractionation, density, and magnetic susceptibility separations, and hand picking of grains. Procedures for such separations related to forensic unconsolidated material examinations are detailed in published references. In most cases, some sample processing is required.
Specimen Preparation Techniques | 5.3.1.3 | Informational Text | Grain size separation can be achieved with clean, laboratory-standard metal sieves or disposable plastic sieve cloth. Alternatively, grain size separation can be achieved by gravitational settling in a fluid following Stokes’ settling law.
Specimen Preparation Techniques | 5.3.1.4 | Recommendation | Segregation of a sample component for SEM/EDS analysis, usually achieved by hand picking particles, is often used to determine the elemental composition of the component to confirm a provisional mineral identification. The morphology of individual grains should be considered when orienting them for optimal imaging and X-ray analysis (see 6.2.3.2).
Specimen Preparation Techniques | 5.3.1.5 | Requirement | When preparing subsamples of particulate material, use procedures that create representative subsamples.
Specimen Preparation Techniques | 5.3.1.6 | Informational Text | Sample splitters, or cone and quartering can be used to create representative subsamples, but both can be impractical for samples of limited quantity.
Specimen Preparation Techniques | 5.3.1.7 | Informational Text | An alternative method of representative sub-sampling appropriate for small quantities of powder, is first to mix the particles, moistening (with an appropriate liquid such as water) to cause particle adhesion, and then selecting one or more sub-samples for analysis from this mixture. NOTE 5—Non-representative subsampling might introduce bias in the modal abundance of particles, affecting downstream analyses.
Specimen Preparation Techniques | 5.3.2 | Section Title | Preparing Unpolished Particles for SEM/EDS:
Specimen Preparation Techniques | 5.3.2.1 | Informational Text | Geological particles can be prepared as unpolished grains to: study their morphology; determine their qualitative elemental composition; or minimize alteration of samples of limited quantity.
Specimen Preparation Techniques | 5.3.2.2 | Informational Text | The choice of mounting substrate is dependent on the analyses. Conductive carbon tabs are a common mounting substrate, but recovering particles from them can be challenging. Particles adhered to polished beryllium or carbon mounts with organic binders (for example, collodion) or conductive paint are easily recovered. NOTE 6—Collodion (nitrocellulose) can be dissolved in acetone.
Specimen Preparation Techniques | 5.3.2.3 | Recommendation | Mineral grains can be mounted either untreated or washed to remove adhering material (for example, grain coatings, decomposition products, explosive residues) prior to analysis. When grains are treated to remove adhering material, this material should be retained for potential subsequent analysis. Grains can be either individually mounted or applied as a dispersion. NOTE 7—Back sieving of sand- or silt-sized grains onto a sticky SEM stub allows grains to be well-spaced on the stub, avoiding the mixing of EDS spectra or blocking of EDS signal that can occur when grains are very closely spaced.
Specimen Preparation Techniques | 5.4 | Section Title | Rocks and Building Materials:
Specimen Preparation Techniques | 5.4.1 | Informational Text | SEM-EDS analysis can be used to examine the phases present and petrographic textures of rocks and building materials (for example, concrete, cinder blocks, bricks, pavers).
Specimen Preparation Techniques | 5.4.2 | Informational Text | The preparation of rocks and building materials for examination can include: direct examination with minimal specimen preparation, polished thin-sections, or embedded and polished (see 5.5).
Specimen Preparation Techniques | 5.4.3 | Informational Text | Direct examination allows for minimal specimen preparation. The major elemental components of samples can be determined using this method of preparation. The elemental composition of surface minerals could be altered from the original source due to chemical weathering.
Specimen Preparation Techniques | 5.4.4 | Informational Text | More commonly, rocks and building materials are analyzed as polished sections (see 5.5).
Specimen Preparation Techniques | 5.4.5 | Informational Text | Aggregate within concrete can be extracted and grains within can be analyzed separately.
Specimen Preparation Techniques | 5.5 | Section Title | Polished Specimens:
Specimen Preparation Techniques | 5.5.1 | Informational Text | Flat specimens are most commonly prepared by cutting, grinding and polishing.
Specimen Preparation Techniques | 5.5.2 | Informational Text | Polished specimens can be prepared from bulk samples, lithic fragments, or individual grains to study grain morphology, internal texture, crystal zoning, weathering rinds, or inclusions.
Specimen Preparation Techniques | 5.5.3 | Informational Text | Loose grains can be embedded in a mounting medium, typically epoxy, and polished.
Specimen Preparation Techniques | 5.5.4 | Informational Text | EDS spectra collected from polished specimens are more reproducible than EDS spectra from unpolished specimens.
Specimen Preparation Techniques | 5.6 | Section Title | Clay-Sized Fraction:
Specimen Preparation Techniques | 5.6.1 | Informational Text | The clay-sized fraction is typically dispersed onto a substrate from a suspension. Dispersion onto a conductive substrate, or application of a conductive coating, will improve SEM imaging and EDS data quality. Due to particle size limitations the elemental analysis of clay-sized particles is commonly performed using the overscanning method (see 6.2.3.3 and 6.2.3.4).
Specimen Preparation Techniques | 5.7 | Section Title | In Situ Analysis of Particles on Items of Evidence:
Specimen Preparation Techniques | 5.7.1 | Informational Text | When isolation of particles from their substrate risks sample loss or alteration, or when the nature of the contact between sample and substrate is of interest, in situ analysis can be conducted (for example, minerals embedded in fabrics, or projectiles).
Specimen Preparation Techniques | 5.7.2 | Informational Text | In some situations, it is necessary to excise portions of the substrate for placement in the SEM chamber.
Specimen Preparation Techniques | 5.7.3 | Requirement | Collect background EDS spectra of the questioned substrate to allow for the differentiation of the EDS signal of the adhering geological materials.
Instrumental Operations and Conditions | 6 | Section Title
Instrumental Operations and Conditions | 6.1 | Section Title | SEM Imaging:
Instrumental Operations and Conditions | 6.1.1 | Section Title | SEM Quality Assurance:
Instrumental Operations and Conditions | 6.1.1.1 | Requirement | Use laboratory-established protocols for instrument operation and maintenance. These protocols are often based on manufacturer’s recommendations.
Instrumental Operations and Conditions | 6.1.2 | Section Title | Selection of Imaging Conditions:
Instrumental Operations and Conditions | 6.1.2.1 | Requirement | Select the analytical conditions for examination by SEM specific to the case and the specimen. The analyst can choose to conduct specific analysis types (see 6.1.4.2 – 6.1.4.5) to meet the goals of the examination.
Instrumental Operations and Conditions | 6.1.2.2 | Requirement | Ensure electron gun and aperture alignment. NOTE 8—Significant changes in accelerating voltage could require adjustments to the aperture alignment.
Instrumental Operations and Conditions | 6.1.2.3 | Requirement | Select an electron imaging detector suitable for observation of the features of interest. NOTE 9—Secondary electron images (SEI) are collected to capture topographic and surface features. Off-axis backscattered electron (BSE) detectors also provide topographic information. NOTE 10—The grayscale intensity value of BSE images is generally proportional to the specimen average atomic number and can be used for assessment of the compositional character of the specimen.
Instrumental Operations and Conditions | 6.1.2.4 | Requirement | Adjust the image brightness and contrast conditions to best observe the specimen and features of interest. Adjust the image brightness and contrast conditions to best observe the specimen and features of interest.
Instrumental Operations and Conditions | 6.1.2.5 | Requirement | Set the accelerating voltage to best observe the sample characteristics under investigation. NOTE 11—Lower voltages provide more surface sensitivity while higher voltages penetrate deeper into the specimen.
Instrumental Operations and Conditions | 6.1.2.6 | Requirement | Optimize the magnification, focus, and stigmation to observe the feature of interest.
Instrumental Operations and Conditions | 6.1.2.7 | Requirement | If sample charging occurs, use a mitigation method as described in 5.2.7 and Note 2.
Instrumental Operations and Conditions | 6.1.3 | Section Title | Image Acquisition:
Instrumental Operations and Conditions | 6.1.3.1 | Requirement | Collect sufficient images to record the features of interest within the specimen.
Instrumental Operations and Conditions | 6.1.4 | Section Title | Assessing Morphological Characteristics:
Instrumental Operations and Conditions | 6.1.4.1 | Requirement | For morphometric assessment, confirm the instrument scale calibration (refer to Practice E766).
Instrumental Operations and Conditions | 6.1.4.2 | Requirement | Particle Surface Texture Examination— Examine and document the surface features of grains; these features are typically compared to reference atlases or to known samples. NOTE 12—Individual particles, typically >250 μm in diameter, are selected for surface texture analysis. To improve imaging, these grain mounts are routinely coated with a conductive material (for example, carbon or gold). NOTE 13—The range of characteristics useful for surface texture analysis is beyond the scope of this guide and can be found in Refs (8) and (9); specific applications of quartz grain surface analysis for forensic purposes can be found in Morgan et al.
Instrumental Operations and Conditions | 6.1.4.3 | Requirement | Particle Shape Examination— Describe the morphology of particles based on: presence and nature of crystal faces, roundness (round to angular); and the relative grain dimensions (for example, equant, elongated, bladed, platy.
Instrumental Operations and Conditions | 6.1.4.4 | Requirement | Particle Size and Sorting— Determine the particle dimensions and distribution of particle sizes (sorting) by SEM imaging. NOTE 14—Particle analysis can use a range of morphometrics and terminology that can be carried out by automated software. NOTE 15—If during specimen preparation the sample was size-fractionated, then any measurements or comparisons of particle size and sorting data should account for this preparation.
Instrumental Operations and Conditions | 6.1.4.5 | Informational Text | The SEM-documented morphology of microscopic particles of biological origin often found in soil (for example, pollen, foraminifera, diatoms) can enable a level of taxonomic identification.
Instrumental Operations and Conditions | 6.2 | Section Title | Elemental Analysis:
Instrumental Operations and Conditions | 6.2.1 | Section Title | EDS Instrument Performance Verification:
Instrumental Operations and Conditions | 6.2.1.1 | Requirement | Perform verification on a regular, documented schedule following instrument-specific requirements and after instrumental maintenance. This verification includes determination of energy calibration, energy resolution, and visual assessment of the Bremsstrahlung background shape. NOTE 16—A standard material (often pure metals such as copper or aluminum) can be used to verify peak positions for both low-energy (~1 keV) and high-energy (~8 keV) peaks.
Instrumental Operations and Conditions | 6.2.2 | Section Title | Selection of EDS Conditions:
Instrumental Operations and Conditions | 6.2.2.1 | Requirement | Use an accelerating voltage sufficient to generate X rays of interest. This value is typically 1.5 to 3 times the energy of the X-ray line for an element of interest. An accelerating voltage of 15 kV to 25 kV is often sufficient for efficient X ray generation. Lower voltages will reduce the analytical volume, increase light element sensitivity, and decrease potential for specimen damage.
Instrumental Operations and Conditions | 6.2.2.2 | Informational Text | When L or M X-ray family lines are provisionally identified, increasing the accelerating voltage can confirm the presence of L or K X-ray family lines of that element.
Instrumental Operations and Conditions | 6.2.2.3 | Requirement | Adjust the beam current, deadtime, and pulse processor time constant to optimize X-ray counts or spectral resolution. NOTE 17—Generally, deadtimes between 10 % to 15 % (for a silicon-drift detector) provide high throughput while minimizing spectral artifacts (for example, sum and escape peaks). NOTE 18—The pulse processor time constant influences collection speed and spectral resolution. Long time constants achieve greater spectral resolution but result in lower count rates.
Instrumental Operations and Conditions | 6.2.2.4 | Informational Text | The optimal working distance for X-ray analysis is instrument dependent. Correctly adjusted, the working distance achieves a take-off angle that optimizes X-ray collection by the EDS detector.
Instrumental Operations and Conditions | 6.2.2.5 | Informational Text | Spectral acquisition times are specimen- and task- dependent. Elements present at low concentrations or light elements (for example, boron, nitrogen, fluorine) typically require longer acquisition times, higher beam current, lower accelerating voltage, or a combination of these parameters to improve detection.
Instrumental Operations and Conditions | 6.2.3 | Section Title | Acquisition of Elemental Composition:
Instrumental Operations and Conditions | 6.2.3.1 | Informational Text | EDS spectra can be collected for a variety of purposes, such as: characterization of individual particles, detection of compositional domains within particles, or over-scanning large areas to survey the elements present in a specimen. NOTE 19—Deflecting the beam from the center of the field of view can result in reduced X-ray counts. This effect is greater at low magnification.
Instrumental Operations and Conditions | 6.2.3.2 | Requirement | Point analysis of a material is achieved by positioning a static beam to a spot of interest. For grains with significant topography, position the static beam on the top of the grain or on a side facing the detector to prevent the specimen from blocking the emitted photon from reaching the detector (see 6.2.4.1).
Instrumental Operations and Conditions | 6.2.3.3 | Informational Text | Bulk analysis of a material is achieved by overscanning a selected region of interest. The degree of homogeneity of the material can be assessed by examining a backscattered electron image to select the size and location of the rastered area. BSE imaging will help to ensure analysis of single compositional domains. When analyzing a material at high accelerating voltages, signal contribution can come from phases/domains at depth.
Instrumental Operations and Conditions | 6.2.3.4 | Requirement | Do not use overscanning (see 3.1.5) for quantitative analysis of large portions of a specimen when there are multiple phases/compositional domains present. NOTE 20—Qualitative overscanning is typically used for detecting elements in fine-grained (clay-sized) preparations. EDS data from the clay-sized fraction can be used together with other methods (for example, X-ray diffraction, FT-IR, or staining) for identification of mineral phases present.
Instrumental Operations and Conditions | 6.2.4 | Section Title | Quality Assessment of Collected EDS Spectra:
Instrumental Operations and Conditions | 6.2.4.1 | Requirement | During and upon completion of the acquisition of an EDS spectrum, assess the data quality for: the shape, intensity, and continuity of the background; peaks; and spectral artifacts (for example, sum peaks, escape peaks). Gaps in the background or sudden breaks in the continuum are an indication of a poor-quality spectrum resulting from specimen geometry (see 6.2.3.2). The low-energy end of the continuum is particularly sensitive to problems. If the low energy continuum is suppressed, this often suggests occlusion of the X-ray signal. Reject spectra of poor quality and recollect.
Instrumental Operations and Conditions | 6.2.4.2 | Requirement | Document artifact peak(s) or re-collect a spectrum under conditions that reduce or eliminate them. Unassigned or ambiguous peaks are noted accordingly.
Instrumental Operations and Conditions | 6.2.4.3 | Informational Text | Once a quality spectrum is collected, X-ray peaks can be assigned to elements by comparison to published tables of elemental X-ray energies or with the assistance of software.
Instrumental Operations and Conditions | 6.2.4.4 | Recommendation | Modern instrument manufacturer software packages provide auto-identification of elements; when using auto-identification, confirm all peak assignments, either by the presence of multiple X-ray lines or checking for the presence of higher energy X-ray family lines, when applicable. If only a single peak with low counts is assigned to an element, the element identification should be considered provisional. Provisional element identifications can be confirmed by collection of replicate spectra, those with a higher number of counts, or by complementary techniques (for example, X-ray fluorescence, or SEM or microprobe equipped with wavelength-dispersive X-ray spectrometers). NOTE 21—When a peak from one element might be occluded by the peak of another element, performing a spectrum fit could be necessary to determine which elements are contributing to a peak. NOTE 22—Due to the energy resolution limits of EDS detectors (approximately 130 eV), there is the potential for overlap of X-ray peaks from multiple elements. Examples of elements encountered in the analysis of geological materials that have overlapping peaks include Pb-S-Mo (molybdenite-MoS2, galena-PbS), Ba-Ti (benitoite-BaTiSi3O9), and Y-P (xenotime-YPO4). For additional examples of peak overlap, see Goldstein et al. Peak interferences (like Pb-S-Mo) are best resolved by peak fitting or spectral synthesis when more than one element could be present. NOTE 23—Element detection using EDS depends on several factors including the matrix in which the element of interest resides, accelerating voltage, beam current, count times, or atomic number. At typical operating conditions (15 kV, ~1 nanoampere of current, and 30 s live time), elements (> fluorine) can be detected at ~0.1 weight %. NOTE 24—Procedures for quality control in element quantitation are beyond the scope of this guide. For guidance on quantitation, see Goldstein et al.
Instrumental Operations and Conditions | 6.2.5 | Section Title | Automated EDS-Based Particle Categorization or Identification:
Instrumental Operations and Conditions | 6.2.5.1 | Informational Text | The automated detection of particles and their elemental compositions can be used to categorize and count particles. The relative abundance of general groups defined by composition can be used in the forensic characterization of geological materials.
Instrumental Operations and Conditions | 6.2.5.2 | Informational Text | Specialized software for automated provisional mineral identification by SEM/EDS has been shown to be useful in examinations of geological material.
Instrumental Operations and Conditions | 6.2.5.3 | Informational Text | Automated particle categorization or identification provides both quantitative modal abundance information as well as morphological assessment of mineral grains.
Instrumental Operations and Conditions | 6.2.5.4 | Informational Text | Procedures for quality control in automated mineral analysis are beyond the scope of this guide.
Interpretation | 7 | Section Title
Interpretation | 7.1 | Informational Text | Reports derived from the forensic analysis of geological materials typically address identification of material, restriction of the possible geographic source area (provenance), and the comparison of two or more materials to determine if they could share a common source. SEM/EDS analysis of any portion of a sample of geological evidence can be included in a report to aid in these three goals.
Interpretation | 7.1.1 | Informational Text | SEM/EDS analysis is usually a component of an overall forensic examination of geological evidence and is not typically used on its own. Use of additional orthogonal methods for forensic examinations is highly recommended. For example, SEM/EDS analysis is commonly conducted following color evaluation (refer to Practice E3254) and detailed PLM examination in the scheme of forensic soil comparisons.
Interpretation | 7.2 | Section Title | Material Identification:
Interpretation | 7.2.1 | Informational Text | In forensic applications, mineral identifications based solely on qualitative EDS analysis are typically considered provisional. Some mineral varieties can be confidently identified by their EDS spectra alone (for example, zircon, ZrSiO4), but for other minerals EDS analysis alone is limited to provisional identification.
Interpretation | 7.2.2 | Recommendation | Provisional mineral identification based on qualitative EDS analysis is performed by visual inspection for the presence, absence, and relative peak area in a spectrum. The identification of a mineral group or species should be supported with the use of mineralogy reference materials, mineral database spectra, or synthesized EDS spectra (for example, DTSA-II). NOTE 25—EDS-based mineral identification and quantification are well-established methods in widespread use in the Earth Sciences. Validation of EDS-based mineral identification is instrument and specimen-specific. Studies have shown strong agreement between X-ray diffraction (XRD) and EDS-based mineral identification for natural and synthetic mineral mixtures. NOTE 26—Minerals within particular groups that display significant solid-solution (for example, feldspar, garnet, pyroxene, amphibole, tourmaline) are often not identifiable by ESD spectra alone, particularly for spectra collected from unpolished specimens. NOTE 27—Some compounds occur as polymorphs (minerals having the same chemical composition but different crystal structures). The identification of a particular polymorph (for example, SiO2, CaCO3, Al2SiO5, TiO2, and KAlSi3O8) is not possible by EDS alone. Differentiation of polymorphs can be achieved using morphology, PLM, XRD, or Raman spectroscopy.
Interpretation | 7.2.3 | Recommendation | Calculations of element abundances using standardless quantitation algorithms common in instrument software packages should be considered estimates.
Interpretation | 7.2.4 | Informational Text | Mineral identifications can be confirmed with the use of an orthogonal instrumental technique (for example, PLM, Raman spectroscopy, XRD, quantitative elemental analysis or by distinctive morphology).
Interpretation | 7.3 | Section Title | Interpreting Morphology:
Interpretation | 7.3.1 | Informational Text | Morphological features including particle surface textures, grain coatings, grain shapes, particle size and sorting, the presence of crystal faces or cleavage, or the identification of microscopic biologically-derived particles (for example, phytoliths, diatoms, foraminifera, pollen) can be used in the interpretation of the geological history of a sample, as a basis of sample comparison, to confirm a provisional mineral identification (see 7.2.4), or to aid in provenance interpretations (see 7.5).
Interpretation | 7.4 | Section Title | Sample Comparisons:
Interpretation | 7.4.1 | Informational Text | The goal of forensic comparisons is to determine whether two samples have exclusionary differences indicating distinct sources. Alternatively, the lack of exclusionary differences supports the proposition that two or more samples could share a common origin.
Interpretation | 7.4.2 | Recommendation | The processes of transfer and persistence of particle assemblages can cause changes in relative proportions of particle types or sizes from the source. X Knowledge of the history of the sample should influence the interpretation of similarities and differences. This can be partially mitigated by comparing similar fractions (size, or grain type).
Interpretation | 7.4.3 | Informational Text | Comparisons Showing Minimal Differences—Two samples of geological material consisting of similar components (for example, morphotypes, elemental types, minerals) in similar relative abundances could have been derived from a common source. If analysis by SEM/EDS supports the conclusion that two samples share similar components (morphological characteristics, provisional mineral identifications, and mineral elemental compositions) in similar relative abundances, then no exclusionary differences are detected.
Interpretation | 7.4.3.1 | Informational Text | EDS spectra of particles in two or more samples can be compared even if the particles’ identities are not known. For particles with similar specimen preparation and data acquisition, spectral overlay can be used for comparisons, where the presence or absence of peaks, peak shapes, and relative intensities are all evaluated as to whether exclusionary differences exist between compared samples.
Interpretation | 7.4.3.2 | Informational Text | The identification of less common features, including minerals, morphologies, mineral compositions, or bulk elemental profiles within both compared samples, increases the strength of association provided by the comparison. For example, soil mineral occurrence data can be used to substantiate the rarity of a given mineral, in general, or specifically within an area of interest. NOTE 28—Minerals that are rare, in general, might be common in the relevant areas of interest to the case. The local characteristics can be ascertained by using more expansive known exemplars, consulting published data, or contacting experts with localized knowledge (for example, regional geological surveys, industry, academic literature, persons with local expertise).
Interpretation | 7.4.4 | Informational Text | Comparisons Showing Differences—If SEM/EDS analysis indicates that there are significant differences with respect to relative abundance, variation, or the omission/addition of particle types (for example, mineral types, grain morphologies, chemical variants of the same mineral), or bulk elemental composition of the fine fraction, evaluate possible explanations for these differences.
Interpretation | 7.4.4.1 | Informational Text | Comparisons Showing Explainable Differences—Explainable differences can include: contamination or alteration of one of the samples due to the effects of fire, stomach acid, mixing or dissolution of soluble mineral phases (for example, anhydrite/gypsum); task-relevant case information; transfer and persistence processes; sample size limitations; and the representativeness of the known exemplars with respect to time of collection or location. If there is no scientifically supported or logical explanation for the observed differences, then these would be considered exclusionary differences.
Interpretation | 7.4.4.2 | Informational Text | Comparisons Showing Exclusionary Differences— Samples with differences detected by SEM/EDS that are not explainable are considered exclusionary differences and indicate that these geological materials were derived from distinct sources.
Interpretation | 7.5 | Section Title | Provenance:
Interpretation | 7.5.1 | Informational Text | The mineralogy, mineral elemental composition, or morphology determined by SEM/EDS can be used to interpret likely and unlikely sources of the material. The methods of this interpretation are beyond the scope of this guide. See Pirrie et al. for a description and methods.
Documentation | 8 | Section Title
Documentation | 8.1 | Informational Text | Refer to Practices E1492 and E620, and ISO 17025 for general guidance on reporting and documentation.
Documentation | 8.2 | Requirement | Document sufficiently to allow a second analyst to understand and evaluate all the work performed, and independently interpret the data.
Documentation | 8.3 | Recommendation | Documentation of geological materials examinations by SEM/EDS should include:
Documentation | 8.3.1 | Recommendation | Specimen preparation procedures (for example, sieving, washing to remove grain coatings, density separation, conductive coating, embedding, and polishing).
Documentation | 8.3.2 | Recommendation | Instrumental conditions used: chamber vacuum conditions, working distance, magnification, beam current/spot size, and accelerating voltage.
Documentation | 8.3.2.1 | Recommendation | For Imaging—Detector, and a scale bar or field-of-view (FOV) dimension.
Documentation | 8.3.2.2 | Recommendation | For EDS Data—Deadtime, time constant, and analytical collection interval (count time or defined total counts).
Documentation | 8.3.2.3 | Recommendation | Relevant controls.
Documentation | 8.3.3 | Recommendation | Information supporting the interpretations of mineral or material identification, provisional mineral or material identification, and references or databases used in identification. Within representative spectra, annotate all peaks associated with elements that are being used for mineral identification.
Documentation | 8.4 | Requirement | For comparisons, document rationale for determining differences that are explainable or exclusionary.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed | Clause involves options; this option was not chosen | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Forensic Document Examination
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics